J Opt Soc Am A Opt Image Sci Vis. Author manuscript; available in PMC 2013 Jun 28.
Published in final edited form as:
J Opt Soc Am A Opt Image Sci Vis. 2013 Apr 1; 30(4): 769–782.
doi: 10.1364/JOSAA.30.000769
PMCID: PMC3695724
NIHMSID: NIHMS448494
PMID: 23595339
# Quantitative orientation-independent DIC microscope with fast switching shear direction and bias modulation
Michael Shribak
Author information Copyright and License information PMC Disclaimer
The publisher's final edited version of this article is available at J Opt Soc Am A Opt Image Sci Vis
Go to:
## Abstract
We describe quantitative orientation-independent differential interference contrast (OI-DIC) microscope, which allows the bias retardation to be modulated and shear directions to be switched rapidly without any mechanically moving. The shear direction is switched by regular liquid crystal cell sandwiched between two standard DIC prisms. Another liquid crystal cell modulates the bias. Techniques for measuring parameters of DIC prisms and calibrating the bias are shown. Two sets of raw DIC images with the orthogonal shear directions are captured within a second. Then the quantitative image of optical path gradient distribution within a thin optical section is computed. The gradient data are used to obtain quantitative distribution of optical path, which represents refractive index gradient or height distribution. Computing enhanced regular DIC images with any desired shear direction is also possible.
Go to:
## 1. Introduction
Differential interference contrast (DIC) light microscope is a polarizing beam-shearing interference system in which the interfering beams are sheared by a small amount, generally by somewhat less than the radius of an Airy disk. The DIC technique produces a monochromatic shadow-cast image that displays the gradient of optical paths. Those regions of the specimen where the optical paths increase along a reference direction appear brighter (or darker), while regions where the path differences decrease appear in reverse contrast. As the gradient of optical path grows steeper, image contrast is significantly increased. Another important feature of the DIC technique is that it produces effective optical sectioning. This is particularly obvious when high numerical aperture (NA) objectives are used together with high NA condenser illumination.
The DIC technique was invented by Fransis H. Smith in 1947 [1,2]. He placed between a pair of the crossed polarizers one Wollaston prism at the front focal plane of the condenser and another Wollaston prism in the back focal plane of objective lens. The splitting angles of the prisms ε1 and ε2 are connected with the focal distances of the condenser and objective fc and fob and the shear amount d by the following relation:
fcε1 = fobε2 = d
(1)
However, in conventional medium- to high-NA objective lenses, the back focal plane is located inside the lens system and therefore not available for insertion of Wollaston prism. The Smith shearing DIC scheme requires special design for the microscope objective lenses so that the Wollaston prism can be incorporated within.
Georges Nomarski, who proposed in 1952 a special polarization prism, took another approach [3,4]. He devised prism, which introduces spatial displacement and angular deviation of orthogonally polarized beams simultaneously. The Nomarski prism can therefore be placed outside of the objective lens.
The DIC microscopy demonstrates remarkable optical sectioning capability, like confocal microscopy. The contrast in DIC is produced by the optical path difference in a small in-focus volume where two interfering beams are spatially separated. Here the beams travel through the different areas of the specimen under investigation. The out-of-focus object introduces practically the same phase disturbance in the both beams because the beams go through almost same area of the specimen. Therefore, the out-of-focus disturbance is suppressed by optical subtraction. It was shown experimentally [5] that the depth of the optical sections, with illumination at 546 nm and 0.95 NA objective lens, could be 0.25 μm for DIC microscopy. Using formulas from a reference [6] one can find that, at the same condition, a theoretical depth of field of conventional brightfield microscope is 0.8 μm. The optical section depth becomes thinner if the shear amount is smaller, and the objective and condenser numerical apertures are larger. The narrow optical sectioning DIC phenomenon is similar to removing an out-of-focus haze in the structured illumination microscopy (SIM) [6]. The SIM employs a single-spatial-frequency grid pattern, which is projected onto the object under investigation. Raw images are taken at three spatial positions of the grid. The out-of-focus picture of the object does not depend on the pattern position. As a result, the out-of-focus haze in SIM is subtracted computationally.
Application of computation subtraction in the DIC should also improve its sectioning capability even further. The computation subtraction of images with different biases is employed in various techniques, such as polarization modulation DIC (PM-DIC) [7,8], differential detection DIC (D-DIC) with polarizing beamsplitter [9], phase-shifting DIC (PS-DIC) [10–12], retardation modulation DIC (RM-DIC) [13–15], and orientation-independent DIC (OI-DIC) [16–18]. In particularly, the PM-DIC removes a background contribution that is insensitive to defocus [8]. It is shown theoretically and confirmed experimentally that a RM-DIC microscope has stronger optical sectioning than a conventional DIC microscope, and the optical section depth is thinner if a Nomarski prism with smaller shear amount is used [15]. Our experiments with the OI-DIC microscope using 100×/1.3NA oil immersion objective lens demonstrated the optical section depth about 0.1 μm. The corresponding field depth of a bright-field microscope would be 0.5 μm.
The regular DIC microscope, however, has shortcomings: it requires the proper orientation of a specimen in relation to the optical system in order to achieve best pictures, and its images are not quantitative. In order to obtain quantitative data it is necessary to capture DIC images with orthogonal shear directions. The previous approaches [11, 19–22] used mechanical rotation either of the specimen or DIC prism. This takes considerable time and the obtained images can be misaligned. Therefore the previous methods in the DIC microscopy were not suitable for imaging live cells. To overcome these limitations we have designed and successfully built a new DIC microscope, which allows the bias and shear directions to be switched rapidly without mechanically rotating the specimen or the prisms [17,18]. In this system we incorporate algorithms that we have developed. The new OI-DIC technique provides quantitative orientation-independent images of phase (dry mass) and phase gradient with highest fidelity and resolution, which cannot be obtained with any other optical microscope.
Go to:
## 2. Principal optical schematics of orientation-independent differential interference contrast (OI-DIC) microscope
A regular DIC microscope produces images, which contrast depends on the beam shear direction and phase shift (bias) between the interfering beams. In order to achieve the highest contrast of a specimen under investigation the researcher should manually rotate a microscope stage to make the shear direction parallel to the specimen’s refractive index gradient, and he/she should optimize the bias by moving a DIC prism laterally or rotating a compensator [6]. The orientation-independent differential interference contrast (OI-DIC) microscope changes the shear direction and the bias quickly and without any mechanical manipulation [17,18]. Instead of a single DIC prism the OI-DIC scheme employs two DIC prisms with liquid crystal variable retarder or liquid crystal 90° polarization rotator in between. This combination creates a beam-shearing assembly, which allows switching the shear direction by 90° at high speed and without the image misalignment. Another liquid crystal variable retarder works as an optical phase shifter that controls the bias. In principle, instead of liquid crystal cells other means for changing the light polarization can be employed.
Principal optical schematics of OI-DIC microscope for the transmitted and reflected light are shown in Fig. 1 and Fig. 2, correspondently. They include light source with bandpath filter, crossed linear polarizer and analyzer, phase shifter, objective lens and image sensor. The transmitted light schematic contains two beam-shearing assemblies and condenser lens. The reflected setup has one beam-shearing assemblies and semi-transparent beamsplitter. The both schematics could also include a displacement compensator. Orientation of the polarizers and birefringent components are shown in the pictures. The phase shifter can be placed anywhere between polarizers and analyzer. It is necessary to keep in mind that the bias will be doubled if the phase shifter is located between beamsplitter and specimen in the reflected light scheme.

Fig. 1
Principal optical schematics of OI-DIC microscope for the transmitted light.

Fig. 2
Principal optical schematics of OI-DIC microscope for the reflected light.
The beam shearing assemblies could employ regular DIC prisms from the microscope. The first prism in assembly is placed in the designed original position and the second DIC prism is displaced along the microscope optical axis by some distance. The large prism displacement causes a strong non-uniformity of intensity distribution across the view filed and reduces the image contrast significantly. One solution of the displacement problem is the using a custom made second DIC prism, which is designed for the specific position. But this way would require knowledge of the DIC prism design and such prism could be very expensive. The second solution is to extract the bare DIC prisms from their original mounts and put them in new holder at close distance. Thickness of each DIC prism is about 1 mm, and thickness of a liquid crystal cell is about 2 mm. A small displacement of the each DIC prism by 2~3 mm from the designed axial position would not deteriorate the image quality substantially. The third solution is to use the second DIC prism in its original mount and a special displacement compensator. The compensator corrects misalignment in the interfering beams caused by displacement of the prism. The axial displacement of the DIC prism changes a transversal distance between two beams and does not change angle between them. The compensator introduces the transversal displacement in position of the extraordinary beam that is equal and opposite to the transversal distance change caused by the DIC prism. Such compensating displacement could be created by a thin calcite plane parallel plate with the optic axis inclined at 45° to the surface, for instance. Detail description of the compensator design is quite cumbersome and it will be done in a future publication.
An example of beam-shearing assembly along with phase shifter is shown in Fig. 3 (center). The picture illustrates the polarization ray tracing in illumination part of the transmitted light microscope (Fig.1). Here we assume that DIC prisms are made of the positive birefringent crystal such as quartz. The diagrams to the left and right in Fig. 3 depict polarization transformations inside the assembly and the corresponding change of shear directions of the output beams. The left column illustrates a case when the beam polarizations between two prisms are preserved and the shear direction is 45° after the second DIC prism. The right column describes a rotation of the output shear direction by 90° caused by switching the beam polarizations after the first DIC prism.

Fig. 3
Design and principle of work of beam-shearing assembly (Color online).
The schematic consists of linear polarizer P, phase shifter, pair of DIC prisms (Nomarski or Wollaston) DIC1 and DIC2 with orthogonal shear directions, and 90° polarization rotator. Shear direction of the first prism DIC1 is chosen as the initial X-direction to describe the orientation of the above mentioned optical elements. Azimuth of the transmission axis of the polarizer P is 45° in order to create equal intensities of polarization components that become spatially divided by the first prism.
The phase shifter introduces bias Γ between the X- and Y- polarization components. Principal axis of the phase shifter is parallel to the shear direction. Dependence of the phase retardation on the applied voltage has to be smooth enough in order to achieve the required bias change accurately and quickly. The initial bias can be adjusted by sliding of the wedge component of one of the DIC prisms. In principle, the phase shifter can be installed with the same orientation in any place in the assembly. A standard untwisted nematic liquid crystal cell, Freedericksz cell, which works in the Electrically Controlled Birefringence (ECB) mode [23–26], can be employed as phase shifter. The Freedericksz cell configuration is different from the twisted nematic configuration typically used in liquid crystal displays. Conventional Freedericksz ECB cell consists of homogeneously parallel aligned liquid crystal directors. The ECB mode uses the applied voltage to change the tilt of the liquid crystal molecules, as a result, the birefringence is changed as a function of the tilt angle. When no voltage is applied, the liquid crystal molecule’s directors are aligned parallel to the cell substrates and the retardance it at a maximum. When a voltage is applied, an electric field is introduced which supplies a torque to the liquid crystal molecules; the retardance is decreased. The molecules near the substrates are not able to fully rotate, and so the retardance does not quite get to zero. Conventional ECB cell has retardance range from 0.1 to two wave (2λ) at 700nm [24,25].
It is also possible to use an ECB cell with vertical alignment (VA ECB) as phase shifter [23,27]. The liquid crystal molecule’s directors are aligned perpendicular (vertical) to the cell substrates if no voltage is applied, and the retardance is at a minimum (zero). Applied voltage turns the liquid crystal directors, and so the retardance is increased. The vertically aligned LC cell exhibits the highest contrast value among all LC cells. Another its merit is fast response speed due to thin cell gap. It does not require a compensation film in order to get zero retardance. The VA ECB has some shortcomings, such as lower retardance range from 0 to half wave (λ/2) at 700 nm, and reduced contrast for off axis rays.
In opposite to the phase shifter, which changes birefringence smoothly in order to obtain the required bias value, the polarization rotator has two states only. In state of OFF the beam passes through the rotator without modifying the X- and Y- polarization components. In state of ON the polarization rotator switches the X- and Y- components. The rotator can utilize conventional ECB or VA ECB liquid crystal cell, which is similar to one used by the phase shifter. In this case the principal axis of the LC cell has to be orientated at 45°. The LC cell switches polarization by 90° (state ON) when the plate retardation is half wave (λ/2), and preserves beam polarization (state OFF) if the plate retardance is 0 or full wave (λ).
It is convenient to utilize a 90° twisted nematic (TN) cell as the rotator [26,27]. Currently the TN cells are commonly used in liquid crystal displays and shutters. They are not expensive and less sensitive to the wavelength variation then ECB cells. In the TN cells the LC anchoring (or rubbing) direction at the input surface is perpendicular to that of the output surface. In the inactivated state (voltage is below threshold), the local LC director undergoes a continuous twist of 90° in the region between the plates. Linearly polarized input light propagates through and, under ideal circumstances, follows the director helix as it propagates through the LC structure. Lets us assume that LC anchoring direction at the input surface is oriented along the X-axis and anchoring direction at the output surface is oriented along the Y-axis. In this case, light with initial X-polarization emerges from the LC cell with Y-polarization, and light with initial Y-polarization comes out with X-polarization. In order to reduce interaction between X- and Y-polarization components during the propagation, the cell’s rubbing direction should be oriented at 0° or 90°. The phase shifter could compensate a possible phase shift between two polarization components introduced by the rotator. When an electric field is applied across the LC layer, the directors become predominantly parallel to the electric field. The twisted structure is thus distorted, and, with voltage above threshold, begins to vanish. The polarization direction of the light is no longer rotated as light passes through the cell.
One of advantages of using the TN cell is that we do not need to know precise dependence of parameters of the cell on the applied voltage. In the transmitted light with use two rotators and can apply the same voltage to the both cell simultaneously using the same generator output. Typically the Freedericksz voltage threshold occurs at 2–3Vrms. In this case, the voltage ~0.5 Vrms will set both rotators in the state ON, and voltage 5 Vrms will put them in the state OFF.
It is promising to employ ferroelectric liquid crystal (FLC) cell in 90° polarization rotator. They are known for their fast switching speed and bistability. The switching speed is very fast, which is usually a few microseconds. The two states are bistable, which means the electric field is not required once the tilting is finished [25,26].
Let us examine in more detail how the beam-shearing assembly works. For example, the initial ray A0 falling on the first prism DIC1 has coordinates (0, 0). Here the first number means a coordinate along the X-axis, and the second one is the coordinate along the Y-axis. The shear plane of the first DIC prism is parallel to the X-axis. The prism deflects a ray with Y-polarization in the negative direction by distance −d/2 and shifts a ray with X-polarization in the positive direction by the same amount. It creates a shear distance d between the components along the X-axis. Thus, the first output ray A1 has X-polarization and coordinates (d/2, 0). The Y-polarized second ray B1 has coordinates (−d/2, 0).
The 90° polarization rotator preserves the linear polarizations of the beams A1 and B1 (state OFF) or turns them by 90° without altering the ray positions (state ON). In the first case (left column) the output rays A2 and B2 are polarized along the X-axis and Y-axis, respectively. In the second case (right column) the rays A′2 and B′2 are polarized along the Y-axis and X-axis, correspondently.
The second prism DIC2 is oriented orthogonal to the direction of the first prism. So, the shear direction of the second prism lies along the Y-axis. The prism moves position of a ray with Y-polarization by distance d/2 in the positive direction and deflects a ray with X-polarization in the negative direction by the same amount.
First we consider a case, which is shown on the left side of Fig. 3, where the rotator state is OFF, and therefore the ray polarizations are not changed. Then ray A2 with X-polarization and coordinates (d/2, 0) passes the second prism with negative displacement −d/2 in the Y-direction, and the output ray A3 has coordinates (d/2, −d/2). In the meanwhile, the coordinates of ray B2 are changed by d/2 in the positive direction along Y-axis, and the coordinates for output ray B3 are (−d/2, d/2). As a result, in the first case the dual beam that falls on the specimen has shear direction −45°. The first case can be described by the following flow-chart:

∣∣∣A0(0,0)(Ex,Ey)∣∣∣→⎧⎩⎨⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪,
DIC1x−shear→∣∣∣A1(d2,0)(Ex,0)∣∣∣→rotatorOFF→∣∣∣A2(d2,0)(Ex,0)∣∣∣→DIC2y−shear→∣∣∣A3(d2,−d2)(Ex,0)∣∣∣
DIC1x−shear→∣∣∣B1(−d2,0)(0,Ey)∣∣∣→rotatorOFF→∣∣∣B2(−d2,0)(0,Ey)∣∣∣→DIC2y−shear→∣∣∣B3(−d2,d2)(0,Ey)∣∣∣
where Ex and Ey are the corresponding amplitudes of polarization components.
For the second case, after the rotator switches the polarizations by 90°, the ray A′2 is linearly polarized along the Y-axis. Therefore the second prism DIC2 changes Y-coordinate of the ray by d/2, and the output ray A′3 has coordinates (d/2, d/2). Also the second prism deflects the X-polarized ray B′2 by −d/2, so that the second output ray B′3 has coordinates (−d/2, −d/2). Hence, the output beam has a shear direction +45°. This case can be illustrated by the next flow-chart:

∣∣∣A0(0,0)(Ex,Ey)∣∣∣→⎧⎩⎨⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪⎪
DIC1x−shear→∣∣∣A1(d2,0)(Ex,0)∣∣∣→rotatorON→∣∣∣A′2(d2,0)(0,Ey)∣∣∣→DIC2y−shear→∣∣∣A′3(d2,d2)(0,Ey)∣∣∣
DIC1x−shear→∣∣∣B1(−d2,0)(0,Ey)∣∣∣→rotatorON→∣∣∣B′2(−d2,0)(Ex,0)∣∣∣→DIC2y−shear→∣∣∣B′3(−d2,−d2)(Ex,0)∣∣∣
As one can see, the both outcome shears have the same amount, which equals to the shear d introduced by a single DIC prism multiplied by 2–√. The shear vectors are mutually orthogonal and oriented at azimuths ±45°. The median axis of two beams passes through point (0, 0) in the both cases. Therefore there is no misalignment between the images.
It is necessary to mention that switch of the light polarization between two DIC prisms affects the outcome bias (Fig. 3). Let’s assume that prisms DIC1 and DIC2 introduce optical path differences Λ1 and Λ2 between beam components polarized parallel and perpendicular to the prism shear plane, respectively. If the rotator state is OFF then polarization of ray A1A3 is parallel to the shear plane of the first DIC prism and perpendicular to the shear plane of the second prism. Thus, in the OFF state the beam-shearing assembly introduces combined bias Λ1−Λ2. In the ON state the rotator switches polarizations by 90°, and polarization of ray A1A′3 is parallel to the shear planes of the both DIC prisms. In this case the combined bias is Λ1+Λ2.
Go to:
## 3. Data processing
### 3.1. Mathematical model of OI-DIC image
A conventional DIC image can be modeled as the superposition of one image over an identical copy that is displaced in the object (specimen) plane by a small shear vector d and phase shifted by bias Λ produced by two DIC prisms, which are located in the illumination and imaging beams correspondently. For simplicity, let’s consider a phase non-birefringent specimen with two-dimensional distribution of optical path φ(x,y). Here we use Cartesian coordinates XOY in the object plane. The specimen is illuminated by monochromatic light with wavelength λ. Intensity distribution I(x,y) in the DIC image depends on the specimen orientation and varies proportionally with the scalar product of shear and optical path gradient vectors:
I(x,y)=I∼sin2{πλ[Λ+d⋅∇φ(x,y)]}+Ic(x,y),
(2)
where Ĩ is the initial beam intensity, ∇φ(x, y) is the optical path gradient vector, and Ic(x,y) corresponds to an offset of the intensity signal, which caused by the stray light.
It follows from the formula (2) that if the shear and optical path gradient vectors are parallel then their scalar product is maximal and the image contrast is maximal also. Where the shear vector is perpendicular to the gradient vector their scalar product is zero and the contrast equals zero correspondently. Thus regular DIC technique shows two-dimensional distribution of the optical path gradient encountered along the shear direction. It is therefore prudent to examine unknown objects at several azimuth orientations.
If X-coordinate is parallel to the shear plane then shear vector d and gradient vector ∇φ(x, y) could be presented in the expanded coordinate forms:
d = (d, 0),
(3)
where d is shear amount, and
(4)
∇φ(x,y)≡(dφ(x,y)dx,dφ(x,y)dy)=(γ(x,y)cosθ(x,y),γ(x,y)sinθ(x,y)),
where γ (x, y) is gradient magnitude and θ (x, y) is gradient azimuth. The gradient magnitude can be interpreted as product of refractive index gradient and optical section depth.
After scalar multiplication of the shear and gradient vectors we obtain the following equation in coordinate form for the regular DIC image:
I(x,y)=I∼sin2{πλ[Λ+d⋅γ(x,y)cosθ(x,y)]}+Ic(x,y).
(5)
As it is shown in Section 2, OI-DIC beam shearing assembly produces two orthogonal shears with shear amount equals to the shear introduced by a single DIC prism multiplied by 2–√. If X-coordinate is parallel to the shear plane of the first DIC prism (see Fig. 3) then shear vectors di could be written in the following coordinate form:
di = (d, (-1)id),
(6)
where i=1, 2 corresponds to the state OFF or ON of 90° polarization rotator (first or second state of shear direction).
Another peculiarity of OI-DIC beam shearing assembly is that the bias introduced by its internal DIC prism is subtracted when the polarization rotator is in OFF state or they are added when the polarization rotator is in ON state. In case of the transmitted light microscope (Fig. 1) the combined bias ΛΣi would be determined by the formula:
Λ∑i = Λ1 + Λ4 + (-1)i(Λ2 + Λ3),
(7)
where Λ1, Λ2, Λ3, and Λ4 are biases contributed by prisms DIC1, DIC2, DIC3, and DIC4, respectively.
The combined bias of the reflected light setup (Fig. 2) is described by similar equation:
Λ∑i = 2Λ1 + (-1)iΛ2.
(8)
Substituting (4) and (6) into (2), and taking into account variable bias Γ created by the phase shifter and combined bias ΛΣi introduced by DIC prisms (see formulae (7, 8), we readily find equation describing OI-DIC microscope:
(9)
Ii(x,y)=I∼sin2{πλ[Γ+Λ∑i+2–√d⋅γ(x,y)cos(θ(x,y)−(−1)iπ4)]}+Ic(x,y),
where note i = 1, 2 corresponds to the 1st or 2nd state of shear direction, −45°or +45° correspondently.
### 3.2. Optimal bias
In order to achieve the best DIC picture it is necessary to apply the optimal bias. For example, we have an extremely shallow depression or elevation that introduces small optical path difference between the studied structure and the media. The high bias would create a bright background and the observed structure would be hardly visible. If bias is low the image background field becomes dark gray and yields very high sensitivity, bringing out image regions with minute phase differences. This conclusion is in agreement with data obtained by Salmon and Tran [28], who indicate that the best contrast is achieved when bias is adjusted to extinguish the light coming from one edge of the object of interest. This occurs when bias equals optical path difference introduced by the edge with the opposite sign. Salmon and Tran found that for the edges of organelles and cells, the optical path difference corresponds to about λ/10th the wavelength or greater, but for microtubules and tiny organelles in cell, optical path difference is very small, less than 1/100th the wavelength of green light. They recommend using about 1/15th–1/20th the wavelength bias for observation of microtubules in order to have sufficient light at the camera. A similar result was found by Schnapp [29].
The bias optimization can be explored with using equation (5). Let consider a simplified sample with a binary gradient magnitude distribution such that a half of the sample has a gradient magnitude of +γ and the other half has a phase difference of −γ. The shear and gradient directions of the sample are parallel (θ = 0°).
We expect the best results to be achieved when DIC images have the highest contrast C:
C=Imax−IminImax+Imin,
(10)
where Imax and Imin are the maximal and minimal intensities in the image.
The maximal and minimal intensities can be determined using equation (5). For the specimen under consideration the contrast of DIC image is the following:
C=sin(2πλΛ)sin(2πλγd)1−cos(2πλΛ)cos(2πλγd)+2IcI∼.
(11)
The ratio Ic /Ĩ is called the extinction ratio. Derivative dC / dΛ of this equation is the next:
dCdΛ=sin(2πλγd)(1+2IcI∼)cos(2πλΛ)−cos(2πλγd)1−cos(2πλΛ)cos(2πλγd)+2IcI∼.
(12)
The derivative is zero at Λ≈γ·d. Thus, the maximal contrast is achieved when bias equals optical path difference in the sample.
According to formula (11) the contrast of microtubules and tiny organelles with phase difference λ/100 equals 0.38 at bias λ/20 and the stray light in the system is absent, 0.27 at bias λ/20 and extinction ratio100, 0.06 at bias λ/4 and both extinctions. The contrast at the small bias is more 6 times the one at the large bias. But it is very important to have a high extinction in order to effectively use a small bias for studying tiny structures. The contrast of cell walls and organelle edges with phase difference λ/10 is 0.79 at bias λ/20 and infinity extinction, 0.72 at bias λ/20 and extinction 100, and 0.59 at bias λ/4 and the both extinctions.
In addition to reducing contrast the low extinction decreases the dynamic range of measurement and can cause a diffraction anomaly in the Airy pattern. The lower extinction is a significant problem in microscopes equipped with high numerical-aperture lenses. To reduce the beam depolarization we can use a polarization rectifier [30,31].
### 3.3 Algorithm for mapping two-dimensional gradient vector field with using low bias swing
The contrast in DIC image is generated in very thin optical section of the specimen. Estimated thickness of the section could be about few tenths of micron in case of employing a high NA objective lens. Additional contrast reduction occurs if the high-resolution DIC prisms with small shear amount are exploited. Taking into account that the maximal difference between refractive indices of cell structures and surrounding water media is 0.2 the maximal optical path difference would be about 0.1μm for the most biological specimens. This number corresponds to 1/5th the wavelength of green light. Actually, in the most cases, the introduced optical path difference is significantly less of that. Therefore it would be beneficial to use a low bias for obtaining the best pictures.
This approach requires preliminary alignment of the OI-DIC system in order to work near the light extinction. At first, it is necessary to set the internal DIC prisms in such position that they mutually compensate or do not introduce the bias, Λ2+Λ3=0 or Λ2=0 (see formulas (7) and (8), respectively). This allows to eliminate the bias that depends on shear direction. Next, using the external DIC prisms and/or the phase shifter, we have to achieve the darkest background in the image. That means the total bias is zero, Γ+ΛΣ=0.
In order to find the two-dimensional distribution of the gradient magnitude and azimuth γ and θ we capture two sets of raw DIC images at shear directions −45° and +45° with negative, zero and positive biases: −Γ0, 0, and +Γ0 [18]. The following group of equations represents these six DIC images:
(13)
Ii,j(x,y)=I∼sin2{πλ[jΓ0+2–√dγ(x,y)cos(θ(x,y)−(−1)iπ4)]}+Ic(x,y),
where j = −1, 0, 1.
Initially two terms are computed (i = 1, 2):
Ai(x,y)=Ii,1(x,y)−Ii,−1(x,y)Ii,1(x,y)+Ii,−1(x,y)−2Ii,0(x,y)tan(πΓ0λ)
(14)
Using equations (13) we can show that
A1(x,y)=tan(22√πλdγ(x,y)cos(θ(x,y)+π4)),A2(x,y)=tan(22√πλdγ(x,y)sin(θ(x,y)+π4)),
(15)
Using the obtained terms we can calculate the quantitative two-dimension distributions of the gradient magnitude and azimuth of optical paths in the specimen as:
γ(x,y)=λ22√πd∑2i=1arctan2[Ai(x,y)]−−−−−−−−−−−−−−−−−√,θ(x,y)=arctan(arctanA2(x,y)arctanA1(x,y))−π4.
(16)
The gradient magnitude represents increment of the optical path difference, which is in nanometers, along lateral coordinate, which is also in nanometers. Thus, the gradient magnitude is unitless. The shear amount d can be measured as it is described in Appendix A or found in Table 1.
### Table 1
Measured shear angles, and computed shear distances and ratios of shear distance to Airy disk radius for various combinations of Olympus DIC prisms and objective lenses at wavelength 532nm
DIC prism type	Shear angle, ε	Shear distance, d (left); and ratio d/rAiry (right)
		UplanFl 10×/0.30	UplanFl 20×/0.50	UplanSApo 30×/1.05 Sil	UplanFl 40×/0.75	UplanSApo 60×/1.20 W	UplanSApo 60×/1.30 Sil	UplanFl 100×/1.30 Oil
U-DICTHR	39 μrad	0.70μm; 0.65	0.35μm; 0.54	0.23μm; 0.76	0.18μm; 0.40	0.12μm; 0.43	0.12μm; 0.47	0.07μm; 0.28
U-DICT	75 μrad	1.35μm; 1.25	0.67μm; 1.04	0.45μm; 1.46	0.34μm; 0.78	0.22μm; 0.83	0.22μm; 0.90	0.14μm; 0.54
U-DICTHC	149 μrad	2.68μm; 2.48	1.34μm; 2.07	0.89μm; 2.89	0.67μm; 1.55	0.45μm; 1.65	0.45μm; 1.79	0.27μm; 1.07

Open in a separate window
Notice that the algorithm considered above employs ratios between intensities of light that have interacted with the specimen. Therefore it suppresses contributions of absorption by the specimen or from nonuniformity of illumination, which can otherwise deteriorate a DIC image.
### 3.4 Algorithm for mapping two-dimensional gradient vector field with using π/4 phase step
If the maximal optical path difference exceeds 1/5th the wavelength we can also employ other algorithms, which are similar to one used in phase-shifting interferometers [32]. For example, let’s consider the modification of algorithm using four frame of intensity data recorded with phase shifts of 0, π/2, π, and 3π/2. This algorithm is common in the interference microscopy [6]. Similar approach we proposed for mapping birefringence at the normal reflection [33].
Under these conditions, the intensity distribution Iij(x,y) on the CCD camera is
(17)
Ii,j(x,y)=I∼sin2{jπ4+2–√πdγ(x,y)λcos(θ(x,y)−(−1)iπ4)}+Ic(x,y).
where j = 0, 1, 2, 3 conforms to four biases, and i = 1, 2 corresponds to states OFF and ON of the polarization rotators. Here we took into account that phase step π/2 is created by bias step λ/4.
Next two terms are calculated:
A∼i(x,y)=Ii,3(x,y)−Ii,1(x,y)Ii,0(x,y)−Ii,2(x,y).
(18)
Using equations (17) we have
A∼1(x,y)=tan(22√πλdγ(x,y)cos(θ(x,y)+π4))A∼2(x,y)=tan(22√πλdγ(x,y)sin(θ(x,y)+π4)).
(19)
Finally we obtain the quantitative two-dimension distributions of the gradient magnitude and azimuth of optical paths:
γ(x,y)=λ22√πd∑2i=1arctan2[A∼i(x,y)]−−−−−−−−−−−−−−−−−−−√,θ(x,y)=arctan(arctanA∼2(x,y)arctanA∼1(x,y))−π4.
(20)
The choice of the algorithms with low bias swing or π/4 step depends on the optical path difference range of the specimen under investigation. The first algorithm is more sensitive for specimen with low optical path difference. It could be optimized to get the best picture by using swing that equals to the maximal optical path difference. The second algorithm does not require preliminary estimation of the optical path difference range. It works well with large optical path difference. The researcher could start with the second algorithm in order to estimate the specimen optical path difference and then apply the first algorithm with optimized swing amount.
### 3.5 Correction of background gradient vector field
When measuring a specimen with small gradient magnitude, it is important to minimize errors that are due to background gradient vector field across the image. Lenses, DIC prisms, LC rotators and phase shifter contribute to the background. The spatially varying background gradient vector field leads to systematic error in the measurement of specimen gradient magnitude and azimuth.
For correction of the background gradient vector field it is possible to employ special procedure, which mathematical description we published early [34, 35]. The procedure is based on specimen images, which are recorded with no structures in the optical path. In order to record the background set of images the object is moved out of the camera view field by use of the x-y translator of the microscope. If empty area around the object is not available one could utilize a defocused picture of the object as the background. The background images are captured at the same beam-shearing assembly settings as for the object itself. Using these images we compute terms Abgi (x, y) or Ãbgi (x, y) according to formulae (14) or (18). Then the differences between the terms Ai (x, y)−Abgi (x, y) or Ãi ( x, y)−Ãbgi (x, y) are used in modified formulas (16) or (20) for determining gradient magnitude and azimuth:
γ(x,y)=λ22√πd∑2i=1arctan2[Ai(x,y)−Abgi(x,y)]−−−−−−−−−−−−−−−−−−−−−−−−−−−√θ(x,y)=arctan(arctan[A2(x,y)−Abg2(x,y)]arctan[A1(x,y)−Abg1(x,y)])−π4,
(21)
or
γ(x,y)=λ22√πd∑2i=1arctan2[A∼1(x,y)−A∼bgi(x,y)]−−−−−−−−−−−−−−−−−−−−−−−−−−−−−√θ(x,y)=arctan(arctan[A∼2(x,y)−A∼bg2(x,y)]arctan[A∼1(x,y)−A∼bg1(x,y)])−π4.
(22)
In addition to correcting background gradient field of the specimen under investigation this procedure also accounts for slight miscalibration of polarization states of the beam-shearing assemblies.
Typically, the background retardance does not change considerably over a time period of many minutes. Therefore the intermediate results Abgi (x, y) or Ãbgi (x, y) can be employed repeatedly to correct, e.g., a time series of images measuring the subtle changes in optical path gradient in a live specimen.
It is possible that intensity of the illumination would fluctuate during the time series image acquisition. We observed such fluctuations using high power mercury arc lamps. Also retardance of the LC phase shifter could slightly vary because of a temperature change. These reasons of an error in the measured specimen gradient vector distribution cannot be corrected with the described background correction procedure. But this small error gradient would be uniform across the image. Therefore if there is an image area with known gradient vector we can use it for additional correction.
For example, the obtained image shows some gradient magnitude γcor and azimuth θcor in an image point where the gradient is expected to be zero. This allows us to compute the correction terms Acori or Ãcori, where i=1, 2, using equations (15) or (19), respectively:
Acor1=A∼cor1=tan(22√πλdγcorcos(θcor+π4)),Acor2=A∼cor2=tan(22√πλdγcorsin(θcor+π4)).
(23)
Then the obtained terms are applied for calculating the corrected gradient magnitude and azimuth for the entire image:
γ(x,y)=λ22√πd∑2i=1arctan2[Ai(x,y)−Abgi(x,y)−Acori]−−−−−−−−−−−−−−−−−−−−−−−−−−−−−−−−−−√θ(x,y)=arctan(arctan[A2(x,y)−Abg2(x,y)−Acor2]arctan[A1(x,y)−Abg1(x,y)−Acor1])−π4.
(24)
or
γ(x,y)=λ22√πd∑2i=1arctan2[A∼i(x,y)−A∼bgi(x,y)−A∼cori]−−−−−−−−−−−−−−−−−−−−−−−−−−−−−−−−−−−√θ(x,y)=arctan(arctan[A∼2(x,y)−A∼bg2(x,y)−A∼cor2]arctan[A∼1(x,y)−A∼bg1(x,y)−A∼cor1])−π4.
(25)
The formulas (24) and (25) allow to receive gradient images, which are corrected on the background retardance, illumination intensity fluctuations, LC retardance variations and systematic errors.
After correction of the gradient vector field, we can compute an enhanced regular DIC image with any shear direction, shear amount and bias retardation using formula (5) with zero offset term. The computed image will be pure phase gradient image with less noise and without deterioration caused by absorption, illumination nonuniformity and light scattering.
### 3.6. Computation of optical path image
The obtained optical path gradient vector field allows us to compute the corresponding two-dimensional scalar field, the optical path, by using numerical integration. The gradient field is a conservative vector field; the computed optical path depends only on the endpoints of that integral, not the particular integration route taken.
Many computation techniques could be employed for determining the optical path, for instance, line integration [16, 36], iterative method [37], various modifications of Fourier integration [11–15, 20, 38, 39], extended Fried algorithm [40], etc. In principle, any of these methods could be employed with us of our OI-DIC microscope. We shall not compare them in this current publication because it expands its size substantially. So it will be done in a separate article.
We use the optical path computation, developed by D. Biggs (KB Imaging Solution LLC, Loomis, CA), which is based on iterative deconvolution approach. This technique employs the same principles as deconvolution techniques normally used to remove out-of-focus haze [41–43].
Go to:
## 4. Experimental verification
The first experimental setup in the transmitted mode, which is shown in Fig.1, was implemented on an upright microscope Olympus BX-61 (Olympus America, Center Valley, PA). The beam-shearing assemblies employed two pairs of regular DIC prisms in their original mounts and 2.2 mm-thick liquid crystal 90° polarization rotators in custom made 7 mm-thick holders. The axial displacement of the outside prism DIC4 in the imaging path was 32 mm up, and the axial displacement of the outside prism DIC1 in the illumination path was 16 mm down. The internal prisms DIC2 and DIC3 were placed in the originally designed positions. We made single displacement compensator, which corrects misalignment in the interfering beams caused by displacements of the both outside prisms. The LC optics was made by Boulder Vision Optics (Boulder, CO). The images were captured by a monochromatic CCD camera Infinity 3-1M (Lumenera, Canada). Using Matlab™ (The MathWorks, Inc., Natick, MA) we developed software for setup control and image processing.
Figure 4 shows OI-DIC images of the hematoxylin-eosin (H&E) stained breast cancer tissue sample, which contains collagenous stroma collagen surrounding the nests of cancer cells. The pictures were taken with using 40×/0.75P UPlanFL objective lens and 546/40nm bandpath filter. The picture size is 186μm ×145μm. We employed two crossed high-resolution DIC prisms U-DICTHR in the top beam-shearing assembly. According to Table 1 the shear distance of single prism equals 0.18μm. Then the shear distance of the crossed prism pair is 0.25μm (0.18μm multiplied by 2–√). The bias modulation amplitude was 0.15λ.

Fig. 4
OI-DIC images of the hematoxylin-eosin (H&E) stained breast cancer tissue sample.
Six raw DIC images Ii,j(x,y) are shown on the left, where i=1 and i=2 correspond to the X- and Y- shear direction, respectively. Notations j=−1, 0, 1 conform to bias −0.15λ, 0 or 0.15λ. All DIC images were captured within 1 sec. In white light the specimen is mostly red. If there is a green filter in the optical path then many specimen structures become dark due to significant absorption in the green-blue spectral region. So the shadow cast DIC features are hardly visible. But they are present, and the computer processing produces quantitative gradient magnitude and azimuth maps (shown in the center). Further computation allows to receive the phase image and regular DIC picture with drastically improved clarity. The enhanced DIC image could be computed with any shear direction and amount without need to rotate either specimen or DIC prisms. The shown enhanced DIC image was computed with the horizontal shear direction, for instance. It could be compared with regular DIC image with the same shear direction, which is displayed on the left top.
Figure 5 shows example of the phase OI-DIC image of a live crane fly spermatocyte during metaphase of meiosis I taken with a Olympus UPlanFl 100×/1.30 oil immersion objective lens and 546/30nm interference filter. Image size is 68 μm × 68 μm. The image acquisition and processing took about 1 second each. The three autosomal bivalent chromosomes are in sharp focus at the spindle equator, along with one of the X-Y sex univalents, which is located on the right. The tubular distribution of mitochondria surrounding the spindle is clearly evident. Both polar flagella in the lower centrosome are in focus, appearing as a letter “L” lying on its side. The experiment was done together with Prof. LaFountain (State University of New York, Buffalo, NY). The phase image was computed by Dr. Biggs (KB Imaging Solution LLC, Loomis, CA) with employing iterative deconvolution approach mentioned above.

Fig. 5
Crane fly spermatocyte (full metaphase of meiosis-I). OI-DIC phase image; brightness is linearly proportional to refractive index.
Go to:
## 5. Conclusion
Here we are proposing new quantitative orientation-independent DIC microscope with high speed switching shear direction and bias modulation. The microscope does not require any mechanically rotating the specimen or the prisms. Therefore there is no image misalignment. All components do not require a special design and are available on the market. We proposed techniques for measuring parameters of DIC prisms and calibrating the bias. Two sets of raw DIC images at the orthogonal shear directions are captured and processed within a second. Then the quantitative image of optical path gradient distribution within a thin optical section is computed. The obtained data are also used to calculate the quantitative distribution of optical phase, which represents refractive index gradient or height distribution. It is possible to generate back the enhanced regular DIC images with any desired shear direction.
Go to:
## Acknowledgments
The author thanks David Biggs of KB Imaging Solution LLC (Loomis, CA) for helpful discussion and providing software for computing phase images and Richard Langill of the Marine Biological Laboratory (Woods Hole, MA) for hardware development. We are also grateful to Shinya Inoué of the MBL, and to James LaFountain of State University of New York (Buffalo, NY) for their encouragement and support. This publication was made possible by Grant Number R01-GM101701 from the National Institute of General Medical Sciences, National Institutes of Health. Its contents are solely the responsibility of the author and do not necessarily represent the official views of the National Institute of General Medical Sciences or the National Institutes of Health.
Go to:
## Appendix A: Measuring shear distance
Shear amount (distance) is the critical parameter of a DIC microscope that determines its contrast, sensitivity, resolution and optical section depth. Another issues with DIC microscopy is that, to derive quantitative information, one must know the amount of image shear. Generally, however, microscope manufacturers do not make that information available. As a result, one must measure this parameter. Munster et al. [38] determined the lateral shift by measuring the distance between the center of the bright spot and the center of the dark spot in image of a submicroscopic transparent latex sphere with bias set at π/2. Mehta and Sheppard [44] measured shear by studying intensity distribution in the back focal plane of the microscope objective lens. Müller et al. [45] used a combination of fluorescence correlation spectroscopy and dynamic light scattering to determine shear. Duncan et al. [46] described a measurement setup with using a standard optical wedge.
The described below technique for determining shear distance is faster and more accurate than mentioned. We measure a shear angle and then compute the shear amount with employing formula (1), as it is explained in the end of this section.
In order to find the shear angle of DIC prism we do not need to know how the passing beam is transformed inside the prism. DIC prism splits an incident monochromatic plane wave into two orthogonally polarized plane wave with slightly different direction of their wavefronts. Fig. 6 illustrates splitting the incident linearly polarized beam into two separate output beams with shear angle ε. Here the shear plane is parallel to the X-axis. As one can see from the picture, the shear angle ε (in radians) equals to derivative of the optical path difference (bias) Λ with respect to the coordinate x:
ε=dΛdx.
(26)

Fig. 6
Angular splitting the incident beam with two orthogonal polarization components Ex and Ey by DIC prism into two output beams with shear angle ε.
The optical path difference Λ is connected with retardance δ, written in degrees, and wavelength λ in the simple way:
Λ=δ360°λ.
(27)
Thus, the shear angle could be found by measuring a derivative of the retardance with respect to the coordinate along the shear direction:
ε=λ360°dδdx.
(28)
Retardance can be determined with a number of various techniques. In particularly, one could employ the return-path compensation method, which we proposed in 1986 [47–49]. This simple technique is similar to the Senarmont compensator [50]. Because of superposition of the probe and reverse beams the return-path arrangement has the following advantages: (i) owing to a double pass of a beam through a specimen under investigation, the sensitivity threshold is decreased by factor of two; (ii) the same optical elements are utilized twice, in the direct and the reverse beams. It makes the measurement set-up less expensive considerably, which is very important, because one high quality optical polarization element could cost several thousands dollars.
Schematic of the return-path setup is shown in Fig. 7. The device operates as follows. A laser radiates the collimated narrow beam, which is unpolarized or circularly polarized. The beam passes sequentially through a small hole in a screen, Glan-Thompson polarization beam splitter, quarter-wave plate QWP, and DIC prism. The DIC prism under investigation is placed on a linear stage with Vernier micrometer. The stage allows to move the prism in the horizontal direction. The beam passes through the quarter-wave plate and DIC prism two times being reflected from a mirror. The beam splitter deflects the orthogonal polarization component toward a screen. The minimum intensity of the deflected laser beam was visually determined by observing the laser spot brightness on the screen. It is also possible to use a photodiode PD to measure the beam intensity. The intensity of the radiation flux of the deflected orthogonal component I′ is described by the following formula [51]:
I′ = I{cosδsin2(ζ-χ)-sinδcos2(ζ-χ)sin2(ζ-ψ)}2,
(29)
where I is the probing beam intensity, δ is retardance of the DIC prism, ψ is angular orientation of its shear plane, χ is orientation of the principal plane of the polarization beam splitter, ζ is orientation of the fast axis of quarter-wave plate.

Fig. 7
Setup for measuring shear angle of DIC prism with employing the return-path birefringence compensator.
The measuring procedure of retardance derivative consists of the following steps. Before placing the DIC prism under investigation in the setup we look this prism between the crossed polarizers. The shear plane would be perpendicular to the observed color fringes or/and the black fringe. Lets note the shear plane orientation.
Next we take out the quarter-wave plate from the setup, and put the DIC prism into the optical path. In this case, beam intensity on the screen is determined by the next expression:
I′ = Isin2δsin22(χ - ψ),
(30)
By rotating the polarization beam splitter from 0° to 180° we achieve two extinctions of the laser spot on the screen. These extinctions differ by 90°. One of the extinction occurs when the beam splitter transmission axis is parallel to the shear plane of the DIC prism, which we noted. In this case the beam splitter angular scale χ points precisely the shear plane orientation ψ.
After that we place the quarter-wave plate QWP into the optical path. The fast axis of the quarter-wave plate has to makes an angle 45° with the shear plane orientation, i.e. ζ = ψ+45°. The beam splitter has to be also rotated by 45°. In this case χ0 = ψ+ 45°. Thus, the quarter-wave plate’s fast axis is parallel to the beam splitter transmission axis. Laser spot on the screen becomes bright. Using formula (29) we could derive the following equation, which describes the spot’s intensity:
I′ = Isin2(δ - 2(χ - χ0)),
(31)
Rotating the beam splitter allows to extinguish the spot on the screen. The DIC prism retardance δ in a point of the laser beam incidence is determined by the simple formula:
δ = 2(χ - χ0),
(32)
where χ is orientation angle of the beam splitter.
Using a linear stage we move the DIC prism in the horizontal direction and measure retardances δ1 and δ2 in two points x1 and x2. The retardance derivative along to the shear direction X could be found by the following formula:
dδdx=δ1−δ2(x1−x2)cosψ=2(χ1−χ2)(x1−x2)cosψ,
(33)
Then we compute the shear angle (in radians):
ε=λ180°(χ1−χ2)(x1−x2)cosψ.
(34)
Instead of finding retardance values in two points only, we can measure dependence of the analyzer angle χ on the laser beam position x and use its derivative in formula above:
ε=λ180°1cosψdχ(x)dx.
(35)
In the described setup we utilized a generic green laser pointer with wavelength 532 nm, beam splitting Glan-Thompson polarizer MSBTS-10-45 (Karl Lambrecht, Chicago, IL, http://www.klccgo.com), and precision achromatic quarter-wave plate AQ-100-0545 (Meadowlark, Frederick, CO, http://www.meadowlark.com). Because beam of our laser pointer has significant linearly polarized component we placed a generic circular polarizer for SLR camera between the pointer and the screen. In particularly we employed Tiffen circular polarizer 52HTCP (www.tiffen.com).
The measurement procedure could be simplified with using a custom built assembly with rotatable polarization beam splitter and quarter-wave plate [48]. This approach would not require taking out and putting in the quarter-wave plate.
Table 1 summaries some results of the shear angle measurements of various DIC prisms currently manufactured by Olympus (Tokyo, Japan, http://www.olympus.com). High-resolution DIC prism U-DICTHR has the smallest shear angle. The prism enables observations with high resolution but less glare even for thick specimen used in developmental and genetic research, such as finely structured diatoms, embryos, zebrafish and C. elegans. All-round prism U-DICT with the intermediate shear angle is suitable for observing a wide range of general specimens, such as tissue. High-contrast DIC prism U-DICTHC has the largest shear angle. Using this prism the high contrast can be obtained even in high magnification observations of thin specimens, such as culture cells.
According to our measurement the shear angle of the all-round DIC prism is double the high-resolution prism and half the high-contrast prism. Therefore some setups with the different objective lenses and the same condenser lens could employ the same condenser DIC prism. The following microscope configurations have interchangeable condenser DIC prisms:
U-DICTHR prism &10× objective, U-DICT prism & 20× objective, U-DICTHC prism & 40× objective;
U-DICT prism & 10× objective, U-DICTHC prism & 20× objective;
U-DICTHR prism & 20× objective, U-DICT prism & 40× objective;
U-DICTHR prism & 30× objective, U-DICT prism & 60× objective;
U-DICT prism & 30× objective, U-DICTHC prism & 60× objective.
For example, condenser DIC prism U-DIC40, which is designed for 40× objective lens and all around DIC prism U-DICT, could be also used with 20× lens and high-resolution DIC prism U-DICTHR. It is interchangeable with condenser prism U-DICT20HR. Condenser prisms U-DIC20 and U-DIC40HC could be also mutually replaced.
We also found that focal distance of the dry top condenser lens is double the oil top condenser lens. Hence, the condenser prism designed for the oil condenser would work the dry condenser and the one step lower resolution objective prism. From another hand, the condenser prism designed for the dry condenser could be paired the oil condenser and the one step higher resolution objective prism. For instance, the dry condenser DIC prism U-DIC60HC, which is designed to work with 60× objective lens and high-contrast prism U-DICTHC, would also work with oil condenser and all around prism U-DICT. So, prism U-DIC60HC is interchangeable with prism U-ODICT60. Then it is also could work instead of U-ODIC30HR, as mentioned early. Another example of interchangeable pair of prisms is U-DIC100 and U-ODIC100HR. The first prism is designed for dry top condenser lens, 100× oil objective and U-DICT prism. The second prism is intended for oil top condenser lens, 100× oil objective and U-DICTHR prism.
The Table 1 shows the corresponding computed shear distance d in the object plane and ratio of the shear distance to the Airy disk radius d/rAiry. For the calculation we used standardized reference focal lengths of the tube lenses Lt for infinity-focused objective lens, which are adopted by several microscope manufacturers.
d=εLtM,
(36)
where M is objective lens magnification. In particularly, the reference focal length Lt is 180 mm for Olympus, 164.5 mm for Zeiss, and 200 mm for Nikon and Leica microscopes [6].
Radius of Airy disk rAiry is determined by the following equation [6]:
rAiry=0.61λNA,
(37)
where NA is objective lens numerical aperture.
The above data is shown for the green light (λ = 532nm). The shear angle was about 4% less for the red light (λ = 641nm).
We have measured parameters of Olympus DIC prism U-DICTH, which was manufactured in previous years. Its shear angle is the same as the prism U-DICTHR. Mehta and Sheppard [44] found the angular shear 74 μrad for U-DICTS prism at wavelength 550nm. This shear corresponds exactly to our result for U-DICT prism. Other our results indicate that shear angles of Nikon 60× I and Zeiss PA63×/1.40III DIC sliders are 76 μrad and 71 μrad, correspondently. According to formula (36) the Nikon slider and 60× objective lens create the shear distance 0.25μm, and the Zeiss slider and 63× objective lens introduce shear 0.18μm.
Go to:
## Appendix B: Calibration of LC variable retarder for bias modulation
Another parameter, which has to be quantified, is a bias amount introduced by the phase shifter. Any method for measuring retardance could be used to find a dependence of bias on the applied voltage. For example, one could employ the return-path compensation method, which we described in Appendix A. However all calibration methods would require taking the phase shifter out of a microscope and placing it into a separate measuring setup. But on the other hand a phase shift of LC variable retarder depends on the temperature, wavelength, set of incidence angles of the passing rays, etc. Therefore it would be convenient to calibrate the LC variable retarder directly in the OI-DIC microscope (Fig. 1 and Fig. 2) in order to avoid mismatch between the calibration and actual data.
For measuring dependence of the bias on the applied voltage we could use preliminary calibrated DIC prims. In a transmitted DIC microscope the total optical path difference (bias) between two interfering beams is created by a pair of DIC prisms. One of them is located in the imaging part, and another in the illumination part. Each of the prisms introduces own bias, which is not uniform and has a gradient. However, the prisms are placed and oriented in that way that they mutually compensate the bias gradient. Thus, the total bias distribution becomes even across the objective back focal plane. The bias can be changed by a lateral shifting of one of the DIC prisms along the bias gradient direction. As it was shown in Appendix A (formula (26)), the unitless bias gradient equals to shear angle (in radians). We can write linear equation for the current bias Λ(x):
Λ(x)=Λ∼+ε(x−x0),
(38)
where x and x0 are the current and initial positions of the prism, Λ˜ is initial bias. The shear angle ε can be found in Table 1 or it can be measured, as it described in Appendix A.
Usually DIC prism is shifted by a translation screw. Then the bias variation could be determined using the pitch of a screw thread p and revolution number R in the following way:
Λ(R)=Λ∼+εpR.
(39)
For example, Olympus high-resolution DIC prism U-DICTHR has translation screw with pitch 2.5 mm. The screw allows maximum 5 rotations. According to the Table 1 the prism shear angle is 40 μrad. Using equation (15) we receive the bias variation Λ = 100 nm per screw 360°-turn and the total range of the bias change Λtot = 500 nm. General Olympus DIC prism U-DICT has 3mm screw pitch and maximum 5 rotations. Consequently its bias variation Λ is 225 nm per screw turn and the total bias range Λtot is 1125 nm.
Let’s consider OI-DIC microscope for the transmitted light (Fig.1), for instance. Before calibration the phase shifter and polarization rotators are removed from the microscope. An empty area of the specimen under investigation is placed in the center of view field. The DIC prisms are adjusted for the maximal image darkness. That means the prisms introduce zero total bias. Then the phase shifter is put into the place.
Because the polarizer and analyzer are crossed and their transmissial axes are oriented at 45° to the principal axes of the orthogonal DIC prisms the beam intensity in the image central area I(Vrms) could be described by the following equation:
I(Vrms)=I∼sin2(πλ(Λ∑+Γ(Vrms)))+Ic,
(40)
where Ĩ is beam intensity after the polarizer, Λ1, Λ2, Λ3, Λ4 and are biases contributed by prisms DIC1, DIC2, DIC3, and DIC4, respectively, ΛΣ = Λ1−Λ2−Λ3+Λ4 is combined bias introduced by all prisms, Γ(Vrms) is bias created by the phase shifter, Vrms is effective (rms) value of the applied voltage, and Ic corresponds to a constant offset of the intensity signal.
If the setup is adjusted for the maximal extinction then biases Λ1, Λ2, Λ3, and Λ4 compensate each other and ΛΣ = 0. Thus, in the initial state of the OI-DIC setup, a dependence of beam intensity I0(Vrms) on the applied voltage is determined by the formula:
I0(Vrms)=I∼sin2(πλΓ(Vrms))+Ic.
(41)
Next we move laterally one of the DIC prisms by the same distance by rotating the translation screw in the clockwise and counterclockwise directions. According to the linear equation (15) in these cases the prism bias will change by amount ±(ε pR). The correspondent dependences of the beam intensity I−1(Vrms) and I1(Vrms) are determined by the following equations:
I−1(Vrms)=I∼sin2(πλ(Γ(Vrms)−εpR))+Ic,
(42)
I1(Vrms)=I∼sin2(πλ(Γ(Vrms)+εpR))+Ic.
(43)
Using the measured beam intensities I0(Vrms), I−1(Vrms) and I1(Vrms) we could obtain the phase shifter bias Γ(Vrms):
(44)
Γ(Vrms)=λ2πarctan[I1(Vrms)−I−1(Vrms)I1(Vrms)+I−1(Vrms)−2I0(Vrms)tan(πλεpR)].
The bias computation formula (44) yields values within [0, λ] interval. In other words, without further analysis, the bias measurement cannot distinguish a certain bias value, say Γ0, Γ0+λ. It is said that the measurement yields wrapped bias information. However, using prior knowledge regarding the continuity of the bias variation, the obtained result can be numerically corrected to cover greater intervals, outside [0, λ], by mathematical operation called phase unwrapping. Mathematically, the unwrapping operation essentially searches for 2π jumps in the signal and corrects them by adding the 2π values back to the signal. For example, Matlab™ has very efficient and fast standard command called “unwrap”.
In order to build the correct bias dependence it is also necessary to take into account a type of LC cell, which is employed in the phase shifter. In particular, retardance of conventional Freedericksz ECB cell decreases with applied voltage. The retardance is minimal and equals about 0.1λ at the highest voltage. In a case of VA ECB cell the retardance is minimal and equals zero when no voltage is applied.
Similar bias calibration method would work for the reflected light OI-DIC setup (Fig. 2). Because each DIC prism is used twice, in the illumination and reflected beams, the bias introduced by lateral movement of the prism is doubled. Therefore formula (44) for the reflected light OI-DIC has to be modified in the following way:
(45)
Γ(Vrms)=λ2πarctan[I1(Vrms)−I−1(Vrms)I1(Vrms)+I−1(Vrms)−2I0(Vrms)tan(2πλεpR)].
Go to:
## References

1. Smith FH. Interference microscope. 2,601,175. US Patent. 1947 Aug 5;
2. Smith FH. Microscopic interferometry. Research (London) 1955;8:385–395. [Google Scholar]
3. Nomarski G. Interferential polarizing device for study of phase object. 2924142. US Patent. 1952 May 14;
4. Allen RD, David GB, Nomarski G. The Zeiss-Nomarski differential equipment for transmitted light microscopy. Zeitschrift für Wissenschaftliche Mikroscopie und Mickroskopische Technik. 1969;69(4):193–221. [PubMed] [Google Scholar]
5. Inoué S. Ultrathin optical sectioning and dynamic volume investigation with conventional light microscopy. In: Stevens J, editor. Three-Dimensional Confocal Microscopy: Volume Investigation of Biological Systems. Academic Press; 1994. pp. 397–419. [Google Scholar]
6. Oldenbourg R, Shribak M. Microscopes. In: Bass M, editor. Handbook of Optics, Third Edition, Volume I: Geometrical and Physical Optics, Polarized Light, Components and Instruments. McGraw-Hill; 2010. pp. 28.1–28.62. [Google Scholar]
7. Holzwarth GM, Webb SC, Kubinski DJ, Allen NS. Improving DIC microcopy with polarization modulation. J Microscopy. 1997;188(Pt 3):249–254. [Google Scholar]
8. Holzwarth GM, Hill DB, McLaughlin EB. Polarization-modulated differential-interference contrast microscopy with a variable retarder. Applied Optics. 2000;39:6288–6294. [PubMed] [Google Scholar]
9. Ooki H, Iwasaki Y, Iwasaki J. Differential interference contrast microscope with differential detection for optimizing image contrast. Applied Optics. 1996;35:2230–2234. [PubMed] [Google Scholar]
10. Hariharan P, Roy M. Achromatic phase-shifting for two-wavelength phase- stepping interferometry. Optics Communication. 1996;126:220–222. [Google Scholar]
11. Arnison MR, Larkin KG, Sheppard CJR, Smith NI, Cogswell CJ. Linear phase imaging using differential interference contrast microscopy. J Microscopy. 2004;214(Pt 1):7–12. [PubMed] [Google Scholar]
12. King SV, Libertun AR, Preza C, Cogswell CJ. Calibration of a phase-shifting DIC microscope for quantitative phase imaging. Proc SPIE. 2007;6443:64430M. [Google Scholar]
13. Ishiwata H, Itoh M, Yatagai T. A new method of three-dimensional measurement by differential interference contrast microscope. Optics Communications. 2006;260:117–126. [Google Scholar]
14. Ishiwata H, Itoh M, Yatagai T. A new analysis for extending the measurement range of the retardation-modulated differential interference contrast (RM-DIC) microscope. Optics Communications. 2008;281:1412–1423. [Google Scholar]
15. Noguchi A, Ishiwata H, Itoh M, Yatagai T. Optical sectioning in differential interference contrast microscopy. Optics Communications. 2009;282:3223–3230. [Google Scholar]
16. Shribak M, Inoué S. Orientation-independent differential interference contrast microscopy. Applied Optics. 2006;45:460–469. [PubMed] [Google Scholar]
17. Shribak M. Orientation-independent differential interference contrast microscopy technique and device. 7,564,618. US Patent. 2003 Dec 17;
18. Shribak M. Orientation-independent differential interference contrast microscopy technique and device. 7,233,434. US Patent. 2003 Dec 17;
19. Fabre L, Inoue Y, Aoki T, Kawakami S. Differential interference contrast microscope using photonic crystals for phase imaging and three-dimensional shape reconstruction. Appl Opt. 2009;48:1347–1357. [PubMed] [Google Scholar]
20. Preza C, Snyder DL, Conchello JA. Theoretical development and experimental evaluation of imaging models for differential-interference-contrast-microscopy. J Opt Soc Am A. 1999;16:2185–2199. [PubMed] [Google Scholar]
21. Hogan H. Getting the small picture. Photonics Spectra. 2003;37(4):58–64. [Google Scholar]
22. Danz R, Gretscher P. C-DIC: a new microscopy method for rational study of phase structures in incident light arrangement. Thin Solid Films. 2004;462–463:257–62. [Google Scholar]
23. Robinson M, Chen J, Sharp G. Polarization Engineering for LCD Projection. Wiley; 2005. [Google Scholar]
24. Chipman RA. Polarimetry. In: Bass M, editor. Handbook of Optics, Third Edition, Volume I: Geometrical and Physical Optics, Polarized Light, Components and Instruments. McGraw-Hill; 2010. pp. 15.1–15.46. [Google Scholar]
25. Scharf T. Polarized Light in Liquid Crystals and Polymers. Wiley; 2007. [Google Scholar]
26. Lueder E. Liquid Crystal Displays: Addressing Schemes and Electro-Optical Effects. Wiley; 2010. [Google Scholar]
27. Gauza S, Wu S-T. Liquid Crystals. In: Bass M, editor. Handbook of Optics, Third Edition, Volume V: Atmospheric Optics, Modulators, Fiber Optics, X-Ray And Neutron Optics. McGraw-Hill; 2010. pp. 8.1–8.40. [Google Scholar]
28. Salmon ED, Tran P. High resolution video-enhanced differential-interference contrast (VE-DIC) light microscopy. Meth Cell Biol. 1998;56:153–184. [PubMed] [Google Scholar]
29. Schnapp BJ. View single microtubules by video light microscopy. Meth Enzymol. 1986;134:561–573. [PubMed] [Google Scholar]
30. Shribak M, Inoué S, Oldenbourg R. Rectifiers for suppressing depolarization caused by differential transmission and phase shift in high NA lenses. Proc SPIE. 2001;4481:163–174. [Google Scholar]
31. Shribak M, Inoué S, Oldenbourg R. Polarization aberrations caused by differential transmission and phase shift in high NA lenses: theory, measurement and rectification. Opt Eng. 2002;41:943–954. [Google Scholar]
32. Hariharan P. Optical Interferometry. 2. Academic Press; London, United Kingdom: 2003. [Google Scholar]
33. Shribak M, Otani Y, Yoshizawa T. Autocollimation polarimeter for measuring two-dimensional distribution of birefringence. Optics and Spectroscopy. 2000;89:155–159. [Google Scholar]
34. Shribak M, Oldenbourg R. Techniques for fast and sensitive measurements of two-dimensional birefringence distributions. Appl Opt. 2003;42:3009–3017. [PubMed] [Google Scholar]
35. Shribak M. Complete polarization state generator with one variable retarder and its application for fast and sensitive measuring of two-dimensional birefringence distribution. J Opt Soc Am A. 2011;28:410–419. [PMC free article] [PubMed] [Google Scholar]
36. Heise B, Sonnleitner A, Klement EP. DIC image reconstruction on large cell scans. Microscopy Research and Technique. 2005;66:312–320. [PubMed] [Google Scholar]
37. Preza C. Rotational-diversity phase estimation from differential-interference-contrastmicroscopy images. J Opt Soc Am A. 2000;17:415–424. [PubMed] [Google Scholar]
38. van Munster EB, van Vliet LJ, Aten JA. Reconstruction of optical pathlength distributions from images obtained by a wide-field differential interference contrast microscope. J Microscopy. 1997;188:149–157. [PubMed] [Google Scholar]
39. Yu SK, Liu TK, Lin SC. Height measurement of transparent objects by adopting differential interference contrast technology. Appl Opt. 2010;49:2588–2596. [Google Scholar]
40. Prahl SA, Dayton A, Juedes K, Sanchez EJ, Lopez RP, Duncan DD. Experimental validation of phase using Nomarski microscopy with an extended Fried algorithm. J Opt Soc Am A. 2012;29:2104–2109. [PubMed] [Google Scholar]
41. Holmes TJ, Bhattacharyya S, Cooper JA, Hanzel D, Krishnamurthi V, Lin W, Roysam B, Szarowski DH, Turner JN. Light microscopic images reconstructed by maximum likelihood deconvolution. In: Pawley JB, editor. Handbook of Biological Confocal Microscopy. Plenum Press; 1995. pp. 389–402. [Google Scholar]
42. Biggs D, Andrews M. Acceleration of iterative image restoration algorithms. Appl Opt. 1997;36:1766–1775. [PubMed] [Google Scholar]
43. Shribak M, LaFountain J, Biggs D, Inoué S. Orientation-independent differential interference contrast microscopy and its combination with an orientation-independent polarization system. Journal of Biomedical Optics. 2008;13:014011. [PMC free article] [PubMed] [Google Scholar]
44. Mehta SB, Sheppard CJR. Sample-less calibration of the differential interference contrast microscope. Applied Optics. 2010;49:2954–2968. [PubMed] [Google Scholar]
45. Müller CB, Weiß K, Richtering W, Loman A, Enderlein J. Calibrating differential interference contrast microscopy with dual-focus fluorescence correlation spectroscopy. Optics Express. 2008;16:4322–4329. [PubMed] [Google Scholar]
46. Duncan DD, Fischer DG, Dayton A, Prahl SA. Quantitative Carré differential interference contrast microscopy to assess phase and amplitude. J Opt Soc Am A. 2011;28:1297–1306. [PubMed] [Google Scholar]
47. Shribak MI. USSR Patent 1,414,097. Device for measuring birefringence of reflecting optical data carrier. 1986 Mar 17;
48. Shribak MI. A compensation method for measuring birefringence. Sov J Opt Technol. 1993;60:546–549. [Google Scholar]
49. Shribak MI. Autocollimating detectors of birefringence. Proc SPIE. 1996;2782:805–813. [Google Scholar]
50. Hartshorne NH, Stuart A. Crystals and the Polarizing Microscope. 4. Edward Arnold; London, United Kingdom: 1970. [Google Scholar]
51. Shribak MI. Polarization separation of the forward and reverse beams in the reading of reflective carriers of information. Sov J Opt Technol. 1986;53:389–391. [Google Scholar]