# Laser-sonic scanner aims to replace mammograms for finding breast cancer
## Imaging an entire breast in 15 seconds
June 18, 2018
Date:
California Institute of Technology
Source:
For women over 40, mammography is a necessary yet annoying procedure to endure every year or two. The technique, while valuable for reducing breast cancer deaths, is less than ideal because it exposes patients to X-ray radiation and requires their breasts to be painfully squished between plates. The plates flatten the breast so the X-rays can more easily pass through it and produce a clear image.
Summary:

Share:
FULL STORY

For women over 40, mammography is a necessary yet annoying procedure to endure every year or two. The technique, while valuable for reducing breast cancer deaths, is less than ideal because it exposes patients to X-ray radiation and requires their breasts to be painfully squished between plates. The plates flatten the breast so the X-rays can more easily pass through it and produce a clear image.

Early detection has been shown to increase breast cancer survival rates, but many women avoid having their mammograms taken as often as they should because of the discomfort involved. A 2013 study found that as many as half of women who were avoiding their mammograms cited pain as the reason why.
Mammography also has trouble with breasts, such as those in young women, that are "radiographically dense," or somewhat opaque to X-rays. And mammography tends to overdiagnose, causing around half of women to receive a false-positive diagnosis at some point in their lives.
Caltech researchers say they have developed something better: a laser-sonic scanner that can find tumors in as little as 15 seconds by shining pulses of light into the breast. The scanning system, known as photoacoustic computed tomography, or PACT, was developed in the lab of Lihong Wang, Caltech's Bren Professor of Medical Engineering and Electrical Engineering.
PACT works by shining a near-infrared laser pulse into the breast tissue. The laser light diffuses through the breast and is absorbed by oxygen-carrying hemoglobin molecules in the patient's red blood cells, causing the molecules to vibrate ultrasonically. Those vibrations travel through the tissue and are picked up by an array of 512 tiny ultrasonic sensors around the skin of the breast. The data from those sensors are used to assemble an image of the breast's internal structures in a process that is similar to ultrasound imaging, though much more precise. PACT can provide a clear view of structures as small as a quarter of a millimeter at a depth of 4 centimeters. Mammograms cannot provide soft-tissue contrast with the level of detail in PACT images, Wang says.
Because the laser light at the currently used wavelength is so strongly absorbed by hemoglobin, PACT can construct images that primarily show the blood vessels present in the tissue being scanned. That's useful for finding cancer because many tumors grow their own blood vessels, surrounding themselves with dense networks of vascular tissue. Those vessels provide the tumors with large amounts of blood and allow the tumors to grow quickly.
During a PACT scan, the patient lies face down on a table that has a recess containing ultrasonic sensors and the laser. One breast at a time is placed in the recess, and the laser shines into it from underneath. Since the scan is quick, taking only 15 seconds, the patient can easily hold their breath while being scanned, and a clearer image can be developed.
"This is the only single-breath-hold technology that gives us high-contrast, high-resolution, 3-D images of the entire breast," Wang says.
The speed with which a PACT scan can be performed gives it advantages over other imaging techniques. For example, magnetic resonance imaging (MRI) scans can take 45 minutes. MRI scans are also expensive and sometimes requires contrast agents to be injected into the patient's blood.
"Gadolinium, for example, is a common contrast agent for MRI which is not totally innocuous," Wang says. "Gadolinium may cause nausea and vomiting, and can remain in the brain for years, with unknown long-term effects. In comparison, PACT is entirely safe; the blood serves as an intrinsic contrast agent, and the laser exposure is well within the safety limits."
In a recent pilot study, the researchers scanned the breasts of eight women using PACT, and correctly identified eight of the nine breast tumors that were present.
The PACT technology has already been licensed by a company, of which Wang is a founder and shareholder, that plans to commercialize it and conduct large-scale clinical studies, Wang says.
"Our goal is to build a dream machine for breast screening, diagnosis, monitoring, and prognosis without any harm to the patient," he says. "We want it to be fast, painless, safe, and inexpensive."
Wang says future research may focus on using PACT for imaging other parts of the body. It could give medical personnel the ability to assess the health of the vascular tissue in the extremities of a patient with diabetes, for example. Diabetes damages blood vessels and, if not properly diagnosed and treated, can cause tissue death, particularly in the feet. PACT is also able to differentiate between oxygenated and non-oxygenated blood; Wang says that could further improve the device's ability to find tumors, since cancer tissue, with its high oxygen demands, deoxygenates blood more rapidly than healthy tissue.

- RELATED TOPICS
  - Health & Medicine
    - Breast Cancer
    - Medical Devices
    - Women's Health
    - Medical Imaging
  - Matter & Energy
    - Medical Technology
    - Optics
    - Detectors
    - Ultrasound
- RELATED TERMS
  - Mammography
  - Breast cancer
  - Breast implant
  - Breast reconstruction
  - X-ray
  - BRCA1
  - Breastfeeding
  - Sunburn
Story Source:
Materials provided by California Institute of Technology. Note: Content may be edited for style and length.

Journal Reference:
Li Lin, Peng Hu, Junhui Shi, Catherine M. Appleton, Konstantin Maslov, Lei Li, Ruiying Zhang, Lihong V. Wang. Single-breath-hold photoacoustic computed tomography of the breast. Nature Communications, 2018; 9 (1) DOI: 10.1038/s41467-018-04576-z

Cite This Page:
- MLA
- APA
- Chicago
California Institute of Technology. "Laser-sonic scanner aims to replace mammograms for finding breast cancer: Imaging an entire breast in 15 seconds." ScienceDaily. ScienceDaily, 18 June 2018. <www.sciencedaily.com/releases/2018/06/180618102645.htm>.

Article
- Open Access
- Published:
# Single-breath-hold photoacoustic computed tomography of the breast
- Li Lin,
- Peng Hu,
- Junhui Shi,
- Catherine M. Appleton,
- Konstantin Maslov,
- Lei Li,
- Ruiying Zhang &
- Lihong V. Wang
Nature Communications  9, Article number: 2352 (2018) Cite this article
volume
- 13k Accesses
- 232 Citations
- 90 Altmetric
- Metrics
  details
## Abstract
We have developed a single-breath-hold photoacoustic computed tomography (SBH-PACT) system to reveal detailed angiographic structures in human breasts. SBH-PACT features a deep penetration depth (4 cm in vivo) with high spatial and temporal resolutions (255 µm in-plane resolution and a 10 Hz 2D frame rate). By scanning the entire breast within a single breath hold (~15 s), a volumetric image can be acquired and subsequently reconstructed utilizing 3D back-projection with negligible breathing-induced motion artifacts. SBH-PACT clearly reveals tumors by observing higher blood vessel densities associated with tumors at high spatial resolution, showing early promise for high sensitivity in radiographically dense breasts. In addition to blood vessel imaging, the high imaging speed enables dynamic studies, such as photoacoustic elastography, which identifies tumors by showing less compliance. We imaged breast cancer patients with breast sizes ranging from B cup to DD cup, and skin pigmentations ranging from light to dark. SBH-PACT identified all the tumors without resorting to ionizing radiation or exogenous contrast, posing no health risks.
## Introduction
Breast cancer is the second most common cancer to affect women in the U.S. and is the second ranked cause of cancer-related deaths. About 1 in 8 (12%) women in the U.S. will develop invasive breast cancer during their lifetime1. Multiple large prospective clinical trials have demonstrated the importance of early detection in improving breast cancer survival2,3,4. While mammography is currently the gold standard used for breast cancer screening, it utilizes ionizing radiation and has lower sensitivity in women with dense breasts5, 6. Ultrasonography has been used as an adjunct to mammography, but suffers from speckle artifacts and low specificity7, 8. Magnetic resonance imaging (MRI) poses a large financial burden and requires the use of intravenous contrast agents that can cause allergy9, kidney damage10, and permanent deposition in the central nervous system11. Diffuse optical tomography has been investigated to provide functional optical contrast. However, the spatial resolution of the current prototypes limits their clinical use12, 13. Overall, each modality has notable advantages and limitations. Photoacoustic computed tomography (PACT) is a promising complementary modality that overcomes many of these limitations.
PACT—ultrasonically imaging optical contrast via the photoacoustic effect—breaks through the ~1 mm optical diffusion limit on penetration for high-resolution optical imaging in deep tissues14, 15. It combines the functional optical contrast of diffuse optical tomography and the high spatial resolution of ultrasonography. The rich contrast bestowed by optical absorption, which is related to various intrinsic and extrinsic contrast origins, enables PACT to perform structural, functional, and molecular imaging16. When a short-pulsed laser irradiates biological tissues, wideband ultrasonic waves (referred to as photoacoustic waves) are induced by transient thermoelastic expansion. The photoacoustic (PA) waves are then simultaneously measured by ultrasonic transducers around the tissue and are used to reconstruct the optical absorption distribution in the tissue17. In the near-infrared (NIR) region, the 1/e attenuation coefficient (1.0–1.3 cm−1)18 for light in an average breast is less than twice that for mammographic X-rays (0.5–0.8 cm−1)19. However, the optical absorption contrast of soft tissue is much higher than X-ray contrast20. For breast imaging, PACT can exploit these advantages to the fullest, offering high spatial and temporal resolutions with sufficiently deep nonionizing optical penetration21, 22. As the principal optical absorber in the NIR region, hemoglobin provides an endogenous contrast for imaging of blood vessels. A high density of blood vessels should correlate with angiogenesis23,24,25, which plays an important role in tumor growth and metastasis26.
Several breast PACT systems have been developed, employing different light illumination and detection schemes27,28,29,30,31,32,33,34,35,36. These systems have advanced PACT toward clinical application, but ongoing limitations remain to be addressed. Here, we consider five main factors: (1) sufficient penetration depth to accommodate most breast sizes and skin colors, (2) high spatial resolution to reveal detailed angiographic structures, (3) high temporal resolution to minimize motion artifacts and enable dynamic or functional studies, (4) minimal limited-view artifacts, and (5) sufficient noise-equivalent sensitivity and contrast-to-noise ratio to detect breast masses.
Specifically, the current systems’ limitations mainly arise from their long scanning times27,28,29 and/or limited-view apertures (i.e., missing data or a <2π steradian solid angle)29,30,31,32,33,34,35. Toi et al. recently reported a photoacoustic imaging system with a hemispherical detector array27, which was modified from a previous design28. Although the design for acoustic detection is slightly different, both used a sparse hemispherical detector array and scanned in a spiral pattern on a plane. The dense sampling, and the nearly isotropic 3D spatial resolutions produced elegant vascular images, but tumor detection was limited by respiratory motion artifacts resulting from the long scanning time (~4 min). Although co-registration partially mitigated the breathing motion distortion, the non-rigidity of the breast compromised the effectiveness. While larger vessels were coregistered, small tumor vessels, which often occur in small clusters, could be challenging to be imaged with partial data and even more difficult to be coregistered. Other groups have used planar transducer arrays30,31,32,33 and arc-shaped arrays34 for breast imaging. However, the limited views of these systems decreased their overall performances37, 38. Consequently, most blood vessels were not well visualized in their images. The same problem occurred with linear transducer arrays, either fixed in position35 or scanned29. A ring-shaped array of 32 elements was developed at presumably relatively low system cost36. However, the low number of elements severely limited the field of view due to the spatial Nyquist sampling criterion, resulting in degradation of image quality39.
Here, we report a significant advancement in breast PACT technology that overcomes all of the aforementioned limitations. Our breast imaging modality—single-breath-hold PACT (SBH-PACT)—is the first PACT system that meets the aforementioned five conditions: (1, 2) Combining 1064-nm light illumination and a 2.25-MHz unfocused ultrasonic transducer array, SBH-PACT achieved up to 4 cm in vivo imaging depth and a 255 µm in-plane resolution (approximately four times finer than that of contrast-enhanced MRI40). (3) Equipped with one-to-one mapped signal amplification and data acquisition (DAQ) circuits, SBH-PACT can obtain an entire 2D cross-sectional breast image with a single laser pulse, or obtain a volumetric 3D image of the entire breast by fast elevational scanning within a single breath-hold (~15 s). The 10 Hz 2D frame rate, currently limited by the laser repetition rate, enables SBH-PACT to observe biological dynamics in a cross-section associated with respiration and heartbeats without motion artifacts. (4) A full-ring 512-element ultrasonic transducer array enables SBH-PACT for full-view fidelity in 2D imaging planes and delivers high image quality. (5) Capitalizing on the optimized illumination method and signal amplification, SBH-PACT achieves sufficient noise-equivalent sensitivity to clearly reveal detailed angiographic structures both inside and outside breast tumors without the use of exogenous contrast agents.
In this pilot study, SBH-PACT was used to image one healthy volunteer and seven breast cancer patients. SBH-PACT clearly identified eight of the nine breast tumors by delineation of angiographic anatomy. These tumors were subsequently verified by ultrasound-guided biopsy. In addition, to improve on the interpretation of images, we developed an algorithm to highlight tumors automatically. Tumors were clearly revealed by SBH-PACT in all breasts even in radiographically dense breasts, which could not be readily imaged by mammography. Taking advantage of the high imaging speed, we demonstrated elastographic SBH-PACT for tumor detection by assessing deformations caused by breathing. SBH-PACT elastography identified the tumor missed in angiographic imaging, and thus improved the sensitivity of tumor detection. At such high spatiotemporal resolutions, SBH-PACT is able to differentiate arteries from veins by detecting blood flow-mediated arterial deformation at the heartbeat frequency.
## Results
### SBH-PACT of healthy breast anatomy and dynamics
The SBH-PACT system is placed underneath a patient bed with minimal separation from the top surface of the bed to the top scanning position of the ultrasonic transducer array (Fig. 1a and Supplementary Fig. 1). With the patient lying prone on the bed, the breast to be imaged is slightly compressed against the chest wall by a soft agar pillow. Compared to craniocaudal or mediolateral breast compression, compression against the chest wall not only avoids pain, but also gives the least thickness breast tissue for light to penetrate from the nipple to the chest wall. The laser illuminates the breast from beneath the patient’s breast, and the ultrasonic transducer array detects photoacoustic waves circumferentially around the breast. The light beam is converted into a donut shape via an axicon lens followed by an engineered diffuser. Compared to a Gaussian beam, the donut beam provides more uniform illumination inside the breast (Supplementary Fig. 2) and also deposits less energy on the nipple and areola, which have a higher concentration of pigment. We take advantage of the low optical attenuation of 1064 nm light to achieve sufficient optical penetration in breast tissue41.
Fig. 1

Representations of the SBH-PACT system. a Perspective cut-away view of the system with data acquisition components removed. b Perspective view of the system with patient bed and optical components removed. DAQ data acquisition system, Pre-amp pre-amplifier circuits
Full size image
To provide 2D panoramic in-plane acoustic detection, we employ a full-ring ultrasonic transducer array with 512 elements. Four sets of 128-channel data acquisition systems provide simultaneous one-to-one mapped associations with the 512-element transducer array (Fig. 1b and Supplementary Fig. 1). Therefore, we acquire photoacoustic signals from a cross section within 100 µs without multiplexing after each laser pulse excitation (Methods). The ultrasonic transducer elements have a central frequency of 2.25 MHz and a one-way bandwidth of more than 95% (Supplementary Fig. 3), providing an experimentally quantified in-plane resolution of 255 µm (Supplementary Fig. 4). The height of each transducer element yields a moderate divergence angle in the elevational direction (~9.0˚ full width at half maximum (FWHM)), yielding a flared diffraction pattern (Supplementary Fig. 5). This pattern enables both 2D imaging of a breast cross section per laser pulse and 3D imaging of the whole breast by scanning elevationally (Methods). Our 3D back-projection algorithm can reconstruct a volumetric image with an elevational resolution of 5.6 mm, which is ~3 times finer than that given by the 2D reconstruction algorithm (Supplementary Fig. 5 and Supplementary Movie 1).
Before imaging breast cancer patients, the performance of SBH-PACT was assessed by imaging a 27-year-old healthy female volunteer. By scanning the transducer array elevationally through her right breast, within one breath hold (~15 s), we revealed the angiographic anatomy from the nipple to the chest wall (Fig. 2a and Supplementary Movie 2). The color-encoded depth-resolved image clearly revealed the detailed angiographic structures of the entire breast (Fig. 2b), visualizing the vasculature down to an apparent vascular diameter of 258 µm (Fig. 2c). To accurately measure the vascular diameters, we identified vessel boundaries in different slices through a correlation-based template matching method42 (Supplementary Fig. 6 and Methods). By doing this, we further investigated the relationship between parent and daughter vessels at vascular bifurcations, which is expressed by the junction exponent (X B )43. We selected a vessel tree in the breast and marked five branch levels with distinct colors (Fig. 2d). At five vascular bifurcations (B1–B5), we calculated the junction exponents as well as the ratios between the cube of the diameter of the parent vessel and the sum of the cubes of the diameters of the daughter vessels (Fig. 2d). For the eight subjects (one healthy volunteer and seven breast cancer patients), we picked five vascular bifurcations in each of their breasts and quantified the average junction exponent (Supplementary Fig. 7), which has a mean value of 2.63 ± 0.34. The junction exponents generally decrease with increasing age44, 45.
Fig. 2

SBH-PACT of healthy breasts. a Vasculature in the right breast of a 27-year-old healthy female volunteer. Images at four depths are shown in increasing depth order from the nipple to the chest wall (also see Supplementary Movie 2). b The same breast image with color-encoded depths. c A close-up view of the region outlined by the magenta dashed box in b, with selected thin vessels and their line spread plots. d A selected vessel tree with five vessel bifurcations, labeled from B1 to B5. At each bifurcation, the diameter relationships between the parent vessel (Dparent) and daughter vessels (Ddaughter) are presented on the right. X B is the junction exponent, and R B is defined as RB=D3parent/(D3daughter_a+D3daughter_b). e Heartbeat-encoded arterial network mapping of a breast cross-sectional image (red = artery, blue = vein). f Amplitude fluctuation in the time domain of the two pixels highlighted by yellow and green dots in e. The pixel value in the artery shows changes associated with arterial pulse propagation (also see Supplementary Movie 3). g Fourier domain of the pixel value fluctuations in f. The oscillation of the arterial pixel value shows the heartbeat frequency at ~1.2 Hz
Full size image
During a breath hold within 10 s, we imaged a cross section of the contralateral healthy breast in one of the breast cancer patients. Working in 2D mode at 10 Hz frame rate, SBH-PACT continuously monitored arterial pulsatile deformation inside the breast (Supplementary Movie 3) by fixing the transducer array at a specific elevational position46. PA signals were analyzed pixel-wise in the frequency domain to identify arteries and veins according to the heartbeat frequency (see Methods) (Fig. 2e). For illustration, we selected a pixel from one artery and one vein (highlighted by round dots 1 and 2 respectively in Fig. 2e) and plotted their pixel value fluctuation (Fig. 2f). The periodic oscillation of the pixel values in the artery indicates that the changes were the result of pulse waves propagating through the arterial network. The oscillation frequency further reveals the subject’s heart rate of ~1.2 Hz (Fig. 2g). Considering that arterial blood has a relatively narrow range of oxygen saturation (sO2)47, average PA signals from arteries can potentially be used to calibrate the local optical fluence (mJ cm−2) deep in the breast, and thus enable accurate quantification of functional parameters (e.g., blood sO2) with an additional laser wavelength (e.g., 750 nm)48, 49.
### SBH-PACT of breast cancer anatomy, segmentation, and elastography
The improved noise-equivalent sensitivity (Supplementary Fig. 8) enabled SBH-PACT to detect breast tumors with fine details, making this imaging modality potentially useful for multiple applications in breast clinical care. We imaged seven breast cancer patients (Fig. 3), with breast sizes ranging from B cup to DD cup (over 99% of the U.S. population has breast sizes of DD cup or smaller50) and skin pigmentations ranging from light to dark (Methods). Angiogenesis, which plays a central role in breast cancer development, invasion, and metastasis, is the essential hallmark by which SBH-PACT differentiates lesions from normal breast tissue23,24,25. Well correlated with the tumor locations shown in mammograms and reported by ultrasound-guided biopsy (Fig. 3a and Methods), SBH-PACT showed eight of the nine tumors by observing higher blood vessel densities associated with tumors in the depth-encoded images (Fig. 3b and Supplementary Fig. 9). We further selected tumor-containing slices perpendicular to the chest wall (marked by white dashed lines in Fig. 3b). In these sagittal (side-view) images, the same tumors, where higher PA amplitude is shown, can be seen at corresponding locations (Fig. 3c). In the X-ray mammograms of Patient 1 (P1) and Patient 6 (P6), the lesions in the dense breasts are barely distinguishable. In comparison, SBH-PACT clearly revealed the tumors not readily seen in mammograms, not withstanding the high radiographical density of the breast.
Fig. 3

SBH-PACT of cancerous breasts. a X-ray mammograms of the affected breasts of seven breast cancer patients. LCC left cranial-caudal, LLM left lateral-medio, LML left mediolateral, LMLO left mediolateral-oblique, RCC right cranial-caudal, RML right medio-lateral. b Depth-encoded angiograms of the eight affected breasts acquired by SBH-PACT. Breast tumors are identified by white circles. For illustration, we marked the nipple of the first patient (P1) with a magenta circle. P1—48-year-old female patient with an invasive lobular carcinoma (grade 1/3); P2—70-year-old female patient with a ductal carcinoma in situ (microinvasion grade 3/3); P3—35-year-old female patient with two invasive ductal carcinomas (grade 3/3); P4—71-year-old female patient with an invasive ductal carcinoma (grade 3/3); P5—49-year-old woman with a stromal fibrosis or fibroadenoma; P6—69-year-old female patient with an invasive ductal carcinoma (grade 2/3); P7—44-year-old female patient with a fibroadenoma in the right breast and an invasive ductal carcinoma (grade 2/3) in the left breast. c Maximum amplitude projection (MAP) images of thick slices in sagittal planes marked by white dashed lines in b. d Automatic tumor detection on vessel density maps. Tumors are identified by green circles. Background images in gray scale are the MAP of vessels deeper than the nipple. e Maps of the relative area change during breathing in the regions outlined by blue dashed boxes in the angiographic images in d. The same tumors are identified by red circles. The elastographic study began with Patient 4, and it revealed all imaged tumors, including the undetected one in d (P7(L))
Full size image
To assist in translation of the technology to a clinical setting, we developed a tumor segmentation algorithm to distinguish tumors automatically. Presumably due to angiogenesis, tumors appear as regions of denser blood vessels in SBH-PACT images (Fig. 3b). To segment tumors automatically, we extracted the vessel skeleton and produced a vessel density map of the breast (local vessel number / local area). The regions with the highest vessel density highlight the breast tumors (Fig. 3d and Methods).
In addition to direct observation of blood vessel density, SBH-PACT detected the difference in compliance between tumors and surrounding normal breast tissue, providing an alternate concurrent contrast to detect breast cancer. Before performing elastography on breast cancer patients, we demonstrated this method on breast-mimicking phantoms (Methods and Supplementary Fig. 10). Working in 2D imaging mode, SBH-PACT quantified the relative area changes in a breast cross section when minor deformations were caused by breathing (Supplementary Movie 4). Because breast tumors are generally less compliant than normal breast tissue51, the regions with lower relative area changes indicated the breast tumor (Fig. 3e). Unlike ultrasonic elastography, SBH-PACT elastography utilized the contrast of hemoglobin and formed area-quantificational grids between vessels (Methods). From only angiographic anatomy detailed by SBH-PACT, the only tumor we missed was located in a marginal region of a D cup breast (P7(L)), where light illumination was insufficient. However, with the addition of SBH-PACT elastography, the missed tumor was identified. Taking advantage of the short time requirement for elastographic measurement (~10 s), SBH-PACT can observe both blood vessel density and tissue compliance simultaneously within ~30 s. Taken together, these two measurements can improve the sensitivity of breast cancer detection.
### Statistics
In this pilot study, SBH-PACT identified eight of the nine biopsy-verified tumors by assessing blood vessel density. Moreover, the initially undetected tumor was subsequently revealed by elastographic SBH-PACT. Pathology reports showed two benign tumors (Patient 5, stromal fibrosis or fibroadenoma; Patient 7, right, fibroadenoma), one ductal carcinoma in situ (DCIS) with a 3/3 nuclear grade (Patient 2), and six invasive carcinomas (all other cases).
Angiogenesis serves as a basis for tumor identification. Considering the diversity among the subjects, we defined high blood vessel densities as values greater than the whole-breast average plus (a) 1.5, (b) 2.0, or (c) 2.5 times the standard deviation, respectively. We calculated and compared the ratios of average vessel density between the high-density region and the normal-density region in each affected and contralateral breast (Supplementary Fig. 11). Receiver operating characteristic (ROC) curves (Fig. 4a) were plotted by varying the threshold of the ratios from 1 to 6. Based on the data from the finite set of subjects, option (b) yielded the largest area (0.90) under the ROC curve. A threshold within (2.26, 2.58) produced a sensitivity (true positive rate) of 88% and a specificity (true negative rate) of 80%. We further performed training and testing studies by obtaining a threshold based on randomly picked six breasts (training set) and then applying the threshold to the remaining seven breasts (testing set). We repeated this procedure ten times and calculated the average sensitivity and specificity (Supplementary Table 1).
Fig. 4

Statistics. a The ROC curves of breast tumor detection based on blood vessel density. σ, standard deviation. b The average vessel density in each tumor and surrounding normal breast tissue. c The relative area change in each tumor and surrounding normal breast tissue caused by breathing. The elastographic study was started with Patient 4. d The longest dimension and center depth of each tumor
Full size image
We then demarcated tumors in each breast and computed the average vessel densities inside and outside the tumors (Methods). The average vessel density ratios between the tumors and the surrounding normal breast tissues were 3.4 ± 0.99 (Fig. 4b). In addition, the mean of the average vessel density ratios of the six malignant tumors was 1.4 times higher than that of the two benign ones (Supplementary Fig. 12).
Since the elastography study began with Patient 4, SBH-PACT elastography identified all five tumors in the subsequent four patients (Fig. 4c). The average breath-induced area change in tumors was around 2 times lower than that in normal breast tissue. As the patient recruitment protocol excluded patients with a mass smaller than 1 cm in diameter in this pilot study, the longest dimension of the smallest tumor we detected was approximately 0.8 cm (Fig. 4d). This tumor was located in the right breast of Patient 7, who was recruited due to a larger tumor in her left breast. However, with 255 µm spatial resolution and refined noise-equivalent sensitivity, SBH-PACT has the potential to detect smaller breast cancers once angiogenesis sufficiently progressed. Patient 3 had DD cup breasts, and her breast was compressed against the chest wall to roughly a cylinder. The tumor in her breast had a depth of ~3.2 cm (elevational distance from the nipple), which was the deepest among the recruited patients (Fig. 4d).
## Discussion
We have developed an SBH-PACT system that integrates deep penetration, high spatiotemporal resolution, sensitive breast cancer detection, and 2D/3D switchable modes. One-to-one mapped low-noise amplifiers and DAQ circuits enabled 2D imaging using a single laser impulse or 3D imaging of an entire breast within a single breath hold (~15 s). The high imaging speed avoided respiration-induced motion artifacts and enabled detection of breast tumors by detailing tumor associated angiogenesis. The donut-shaped optical illumination and panoramic acoustic detection provided a more uniform fluence distribution in deep tissue and best in-plane coverage of ultrasound reception, respectively, delivering high image quality. Furthermore, considering the low cancer detection rate (0.41%)52, even though modern mammography uses a low dose of ionizing radiation, the risk-to-benefit ratio (e.g., 8–17% https://www.ncbi.nlm.nih.gov/pubmed/11388061 for 40–50 year-old women)53, 54 is considered high. In comparison, SBH-PACT requires neither ionizing radiation nor an exogenous contrast agent, yielding zero risk.
In SBH-PACT, the laser beam was broadened into a donut shape with an outer diameter of ~10 cm, depositing light with an average laser fluence of ~20 mJ cm−2 on the breast surface (only 1/5 of the American National Standards Institutes safety limit)55. This outer radius covered most breasts and provided satisfactory SNR in breast images. Merely assessing blood vessel density, we missed one tumor, which was located in an insufficiently illuminated marginal region of a D cup breast (P7(L) in Fig. 3a). Therefore, SBH-PACT can potentially improve sensitivity further in breast cancer detection if equipped with a more energetic laser, which will allow us to enlarge the illumination area and increase the optical fluence.
Our automatic tumor segmentation algorithm made it easier to recognize tumors by highlighting the suspicious affected region with the highest vessel density. In addition, the high 2D imaging speed of SBH-PACT (10 Hz frame rate) enabled us to perform elastographic measurements and to improve on breast cancer detection. The capability of SBH-PACT to map arterial distribution can potentially be useful in diagnosing artery-related diseases56,57,58. In addition, the knowledge of vessel diameters and average PA signals from arteries can be used to calibrate the local optical fluence48, 49, thus providing accurate spectral sO2 measurement in deep tissue.
By developing this advanced breast imaging modality, we provided a promising tool for future clinical use including not only screening, but also diagnostic studies to determine extent of disease, to assist in surgical treatment planning, and to assess responses to neoadjuvant chemotherapy. Compared to mammography, SBH-PACT utilizes non-ionizing radiation, shows early promise for sensitivity in radiographically dense breasts, and imposes less or no pain by only slightly compressing the breast against the chest wall. Because the average hemoglobin concentration in malignant tumors is generally twice that in benign tumors59,60,61, SBH-PACT has the potential to distinguish malignant tumors from benign tumors by quantifying blood vessel densities in the tumor (Supplementary Fig. 12). Based on the vessel density in the two benign tumors and the six detected malignant ones in this study, the threshold of the vessel density ratio between tumors (either malignant or benign) and healthy tissues should be set within the range of (2.72, 2.76) to optimally differentiate malignant tumors from benign ones; however, more patients need to be imaged to achieve statistical significance. Using hemoglobin as the contrast, SBH-PACT can potentially monitor breast cancer’s response to neoadjuvant chemotherapy by acquiring information similar to that of contrast-enhanced MRI, yet with finer spatial resolution, higher imaging speed, and only endogenous contrast.
## Methods
### System construction
The SBH-PACT system mainly comprises of an illumination laser, an ultrasonic transducer array, signal amplification/acquisition modules, a linear scanning stage, and a patient bed (Fig. 1). The 1064-nm laser beam (PRO-350-10, Quanta-Ray, 10-Hz pulse repetition rate, 8–12-ns pulse width) was first passed through a lab-polished axicon lens (25 mm diameter, 160° apex angle), then expanded by an engineered diffuser (EDC-10-A-2 s, RPC Photonics) to form a donut-shaped light beam. The laser fluence (20 mJ cm−2) was within the American National Standards Institutes (ANSI) safety limit for laser exposure (100 mJ cm−2 at 1064 nm at a 10-Hz pulse repetition rate)55. To synchronize the SBH-PACT system, the laser’s external trigger was used to trigger both the data acquisition systems and the linear scanner.
To achieve 2D panoramic acoustic detection, we employed a 512-element full-ring ultrasonic transducer array (Imasonic, Inc.; 220 mm ring diameter; 2.25 MHz central frequency; more than 95% one-way bandwidth; Supplementary Fig. 3). Each element had a flat-rectangular aperture (5 mm element elevation size; 1.35 mm pitch; 0.7 mm inter-element spacing). The ultrasonic transducer array housing was mounted on a stainless steel rod (25 mm diameter) and enclosed in an acrylic water tank. A linear stage (THK America, Inc., KR4610D) was fixed beneath the water tank and moved the transducer array elevationally via the stainless steel rod. Four sets of lab-made 128-channel preamplifiers (26 dB gain) were placed around the water tank, connected to the ultrasonic array housing via signal cable bundles. Each set of preamplifiers was further connected to a 128-channel data acquisition system (SonixDAQ, Ultrasonix Medical ULC; 40 MHz sampling rate; 12 bit dynamic range) with programmable amplification up to 51 dB. The digitized radio frequency data were first stored in an onboard buffer, and then transferred to a computer through a universal serial bus 2.0 (Supplementary Fig. 1). The data acquisition systems were set to record PA signals within 100 µs after each laser pulse excitation.
The patient is positioned prone with one breast dependent and placed into a large aperture in the bed. An agar pillow affixed on top of an acrylic tube lightly pressed the breast against the chest wall. The bed top was covered by cushioning memory foam. The water tank was fully filled with water preheated to a temperature of 35 °C. Both the patient bed and the SBH-PACT system were supported by T-slotted aluminum frames.
### Half-time PA reconstruction in 2D and 3D modes
We used the half-time universal back-projection (UBP) algorithm62 to reconstruct all images in this work. In 2D imaging mode, the time-domain PA signals generated by each laser pulse were back-projected to a 2D imaging plane. Determined by the acoustic divergence angle (~9.0˚) at FWHM in the elevational direction (Supplementary Fig. 5), the elevational resolution at the center was ~16.1 mm.
Alternatively, when working in 3D mode, the ultrasonic transducer array scanned the entire breast from the chest wall to the nipple. The time-domain PA signals acquired at all elevational scanning steps were then back-projected simultaneously into the 3D space. To accommodate the acoustic divergence angle in the elevational direction, 3D-UBP added a weight to the back-projected PA signals at different elevational divergence angles (Supplementary Fig. 5). To accurately reconstruct objects in the Fraunhofer zone, we back-projected PA signals from virtual transducers located at the transition points between the Fresnel and Fraunhofer zones63. Sharing the same in-plane resolution as the 2D mode, 3D-UBP provided an improved elevational resolution of 5.6 mm.
The full-ring transducer array with 512 elements could spatially well sample objects — according to the spatial Nyquist criterion — within a field of view (FOV) of ~39 mm in diameter39. To eliminate aliasing caused by under-sampling in regions outside of this FOV, we low-pass filtered PA signals with cut-off frequencies determined by the distance to the center of the ring array.
Each volumetric image was first reconstructed with a voxel size of 1 mm in the elevational direction and 0.1 × 0.1 mm2 on the horizontal plane. We then batch-processed all the reconstructed images to improve contrast (batch processing before vesselness filtering in Supplementary Fig. 9). In each horizontal slice, we applied Hessian-based Frangi vesselness filtration39 to enhance the contrast of blood vessels with diameters ranging from 3 to 12 pixels. In each filtered slice, adaptive thresholding was used to segment blood vessels42, followed by morphology filtration for single-pixel elimination. In the elevational direction of each filtered volumetric image, we selected voxels with the largest PA amplitudes and then projected their depths to form a 2D image. We applied median filtration with a window size of 3 × 3 pixels to the depth image. Another median filtration with a window size of 6 × 6 pixels was further applied inside the segmented vessels to the segmented vessels’ depths. Different RGB (red, green, blue) color values were assigned to discrete depths (vertical color bar in Fig. 2b). Finally, the 2D depth-resolved color-encoded image was multiplied by the MAP image pixel by pixel to represent the maximum amplitudes (horizontal color bar in Fig. 2b). To further reduce noise and improve image quality, we also tuned the above parameters in 2D slices at different depths (custom processing in Supplementary Fig. 9). As shown in Supplementary Fig. 9, the structures in all three sets of images match well with each other, showing the fidelity of the vesselness filtering and custom processing.
### Vascular diameter measurement
Vascular diameters were accurately measured by identifying vessel boundaries through a correlation-based template matching method42. The templates were generated through simulation (Supplementary Fig. 6). The impulse responses of all ultrasonic transducers were used to simulate the images of vessels with different sizes (0.5–2.0 mm) and orientations. The diameters of vessels chosen from the SBH-PACT breast images were quantified by matching the reconstructed vessel images with the generated templates.
### Arterial vessel mapping
Working in 2D imaging mode, SBH-PACT was able to monitor blood flow-mediated arterial fluctuation (Supplementary Movie 3). After removing displacement through rigid transformation, we analyzed the pixel value fluctuation during a patient’s breath hold (~10 s). We found that arteries fluctuated much more than veins at the frequency of the heartbeat. The fluctuation of the pixel values in the artery indicated the changes associated with arterial pulse propagation (Fig. 2g).
To separate fluctuations caused only by heart beats, frames with strong motion caused by body movement were first removed. The entire imaging field was then divided into 16 slightly overlapping subdomains. In each subdomain, we chose the first frame as the reference frame; other frames were registered to it through rigid transformation, optimizing the frame–frame correlation. In each subdomain, a Gaussian filter with a radius of 0.2 mm was applied to all registered frames to reduce high spatial-frequency noise. We then applied Fourier transformation to each pixel’s value through all the frames. The fluctuations in pixel values induced by arterial pulse propagation were quantified within the frequency range (1.0–1.6 Hz) of heartbeat cycles64.
### Tumor segmentation
SBH-PACT showed breast masses by revealing a greater density of blood vessels, presumably due to angiogenesis, in tumor regions. To segment tumors automatically, we extracted the vessel skeleton and produced a vessel density (number of vessels / area) map of the breast. The regions with the highest vessel density highlighted the breast mass of interest (Fig. 3d).
The dense vessels in the nipple would affect the automatic tumor segmentation. Therefore, the shallowest slices containing the nipple were first removed. The remaining slices were used to generate the MAP image. A vessel mask was generated from the MAP by Hessian filtering and threshold-based segmentation. Based on the mask, vessel centerlines were extracted by removing boundary pixels. The vessel centerlines were broken into independent vessels at junction points. To reduce noise further, we removed independent vessels with lengths less than 3 pixels (255 µm spatial resolution divided by 100 µm pixel size is approximately 3). A 2 mm × 2 mm window was then used to scan the entire image. At each scanning location, the number of vessels (independent segments) inside the window was counted and assigned to the center pixel in the window. The vessel density was quantified as the number of vessels divided by the window area. To compute the average vessel density of the whole breast, we included pixels inside a 10 cm-diameter circle around the image center.
To demarcate breast tumors from MAP images, we first identified suspicious regions where blood vessel densities were higher than a threshold, which was set to each whole-breast’s average plus 2.0 times the standard deviation. Among the eight affected breasts, the smallest suspicious region had a diameter of 1 mm. We then counted the numbers of pixels in each contiguous region and rejected the regions with pixel counts fewer than 1855 (18.55 mm2) to eliminate false positive cases (Supplementary Fig. 13). The remaining contiguous regions were labeled as tumors and the smallest one had a longest dimension of 8 mm (Fig. 4d). In comparison, contrast-enhanced MRI on a 1.5 Tesla scanner can detect breast tumors as small as 4 mm, which is similar to the smallest size of tumors detectable by X-ray mammography65, 66.
### Elastographic study
SBH-PACT’s high imaging speed enabled differentiation in compliance between tumors and surrounding normal breast tissues, providing another contrast for detecting breast cancer. We first performed SBH-PACT elastographic measurements on a breast phantom. The phantom comprised a ball with 7% agar (mimicking breast tumor) embedded in a base of 2% agar (mimicking normal breast tissue)67. Chopped human hair was uniformly distributed in the phantom to mimic small blood vessels. Working in 2D imaging mode, SBH-PACT quantified the relative area changes in a cross section when minor deformations were induced by periodic compressions (~0.25 Hz) on top of the phantom. Due to the low elevational sectioning power of 2D imaging, objects in 2D frames were mainly influenced by coronal dilation instead of elevational displacement. Accordingly, SBH-PACT elastography clearly revealed the agar ball with correct size and location (Supplementary Fig. 10). No obvious differences were observed in the concentration of the hair fiber between the balls and the phantom base.
To assess deformations over time, the first frame was taken as a reference. Other frames were registered to the first frame through a non-rigid demon algorithm68 in Matlab. For each pixel of registered frames, the standard deviation (STD) of the value variations was calculated. Pixels with relatively small STDs were stably registered and were used for deformation quantification. The entire image was then segmented into 2 mm × 2 mm squares. One stably registered pixel was chosen from each square, and triangular grids were further generated from these registered pixels. The triangular grids were mapped back to the original unregistered frames, and their areas were calculated. For each grid, Fourier transformation was applied to quantify the area variation at the frequency of periodic compression, and amplitudes were assigned to the pixels inside this triangle to generate the deformation map. To further reduce noise, 100 deformation maps were generated with randomly registered pixels in the squares. The final image is the average of the 100 deformation maps.
To conduct SHB-PACT elastography of the breast, patients were asked to breathe normally. The chest wall pushed the breast against the agar pillow, elevationally generating a deformation of the breast in the coronal plane. We used the same method to quantify the change of area between blood vessels in the breast. Tumors, being stiffer, could be identified in areas with less deformation than normal breast tissue.
### Patient recruitment and human experiment protocols
All the human experiments followed protocols approved by the Institutional Review Board (IRB) and Protocol Review and Monitoring Committee (PRMC) of Washington University in St. Louis. We enrolled one healthy volunteer and seven female patients with consent documents signed. All recruited patients met the following inclusion criteria: (1) Patients were newly diagnosed with breast tumors highly suspicious for malignancy, larger than 1 cm in diameter, and were eligible for percutaneous biopsy; (2) Patients were females >18 years of age; (3) Patients were able to understand and willing to sign a written informed consent document. The exclusion criteria included (1) patients whose weight exceeded 300 lbs (the weight limit of the steps for mounting the bed); (2) patients who were pregnant; (3) patients with uncontrolled intercurrent illness including, but not limited to, ongoing or active infection of the breast and/or axilla, symptomatic congestive heart failure, unstable angina pectoris, cardiac arrhythmia, or psychiatric illness/social situations that would limit compliance with study requirements.
Photoacoustic imaging was performed after a standard of care (SOC) work-up, but in advance of percutaneous biopsy. This order of events was designed to minimize confounding imaging findings related to biopsy-induced hemorrhage. Patients underwent only one PACT imaging study, which took less than 10 min. We imaged both the contralateral and affected breasts. For the abnormal breast, we analyzed the tumor size, tumor depth, blood vessel density, and signal amplitude in the breast images. Our analysis of tumor size/depth was further compared with the standard imaging results (mammography and ultrasonography). To identify the tumor types and grades, histopathology results from the SOC biopsy were used as the ground truth for interpretation of the results.
### Standard of care work-up, percutaneous biopsy, and pathologic diagnosis
Patients were imaged at the Joanne Knight Breast Health Center at Siteman Cancer Center in partnership with Washington University School of Medicine (WUSM), Mallinckrodt Institute of Radiology. All standard of care imaging was performed by subspecialist breast imaging physicians who are faculty of WUSM. Using established clinical protocols, abnormalities were identified either through routine screening mammography, or diagnostic evaluation in symptomatic patients. Pre-biopsy work up included combinations of digital mammography, digital breast tomosynthesis, and ultrasound. Formal BI-RADS (breast imaging, reporting and data system) assessments were assigned in all cases, with appropriate recommendation for biopsy. Eligible patients were approached and informed consent obtained in advance of the biopsy. Image-guided percutaneous biopsy was obtained using real-time ultrasound guidance and a 12-guage or 14-guage spring-loaded biopsy needle (chosen at the discretion of the performing physician.). Core specimens were submitted in formalin to the pathology department for histologic analysis as per normal routine at the institution. All cases were reviewed following receipt of the final pathology report to determine radiologic-pathologic correlation. Some patients underwent contrast enhanced breast MRI following confirmation of malignancy.
### Data availability
All data are available within the Article and Supplementary Files, or available from the authors upon request.

## References
1. Siegel, R. L., Miller, K. D. & Jemal, A. Cancer statistics, 2017. CA Cancer J. Clin. 67, 7–30 (2017).
  Article PubMed Google Scholar
2. Dizon, D. S. et al. Clinical cancer advances 2016: annual report on progress against cancer from the American Society of Clinical Oncology. J. Clin. Oncol. 34, 987–1011 (2016).
  Article PubMed PubMed Central CAS Google Scholar
3. Miller, A. B. et al. Twenty five year follow-up for breast cancer incidence and mortality of the Canadian National Breast Screening Study: randomised screening trial. BMJ 348, g366 (2014).
  Article PubMed PubMed Central Google Scholar
4. Burton, R. & Bell, R. The global chanllenge of reducing breast cancer mortality. Oncologist 18, 1200–1201 (2013).
  Article PubMed PubMed Central Google Scholar
5. Pinsky, R. W. & Helvie, M. A. Mammographic breast density: effect on imaging and breast cancer risk. J. Natl Compr. Canc. Netw. 8, 1157–1165 (2010).
  Article PubMed Google Scholar
6. Freer, P. E. Mammographic breast density: impact on breast cancer risk and implications for screening. Breast Imag. 35, e352140106 (2014).
  Google Scholar
7. Devolli-Disha, E., Manxhuka-Kerliu, S., Ymeri, H. & Kutllovci, A. Comparative accuracy of mammography and ultrasound in women with breast symptoms according to age and breast density. Bosn. J. Basic. Med. Sci. 9, 131–136 (2009).
  Article PubMed PubMed Central Google Scholar
8. Hooley, R. J., Scoutt, L. M. & Philpotts, L. E. Breast ultrasonography: state of the art. Radiology 268, e13121606 (2013).
  Article Google Scholar
9. Murphy, K. J., Brunberg, J. A. & Cohan, R. H. Adverse reactions to gadolinium contrast media: a review of 36 cases. Am. J. Roentgenol. 167, 847–849 (1996).
  Article CAS Google Scholar
10. Perazella, M. A. Gadolinium-contrast toxicity in patients with kidney disease: nephrotoxicity and nephrogenic systemic fibrosis. Curr. Drug Saf. 3, 67–75 (2008).
  Article PubMed CAS Google Scholar
11. Ibrahim, D., Froberg, B., Wolf, A. & Rusyniak, D. E. Heavy metal poisoning: clinical presentations and pathophysiology. Clin. Lab. Med. 26, 67–97 (2006).
  Article PubMed Google Scholar
12. Choe, R. et al. Diffuse optical tomography of breast cancer during neoadjuvant chemotherapy: a case study with comparison to MRI. Med. Phys. 32, 1128–1139 (2005).
  Article PubMed Google Scholar
13. Culver, J. P. et al. Three-dimensional diffuse optical tomography in the parallel plane transmission geometry: evaluation of a hybrid frequency domain/continuous wave clinical system for breast imaging. Med. Phys. 30, 235–247 (2003).
  Article PubMed CAS Google Scholar
14. Xia, J., Yao, J. & Wang, L. V. Photoacoustic tomography: principles and advances. Electromagn. Waves 147, 1–22 (2015).
  Article Google Scholar
15. Razansky, D. et al. Multispectral opto-acoustic tomography of deep-seated fluorescent proteins in vivo. Nat. Photon. 3, 412–417 (2009).
  Article ADS CAS Google Scholar
16. Yao, J., Xia, J. & Wang, L. V. Multi-scale functional and molecular photoacoustic tomography. Ultrason. Imag. 38, 44–62 (2016).
  Article Google Scholar
17. Zhou, Y., Yao, J. & Wang, L. V. Tutorial on photoacoustic tomography. J. Biomed. Opt. 21, 061007 (2016).
  Article ADS PubMed Central Google Scholar
18. Durduran, T. Bulk optical properties of healthy female breast tissue. Phys. Med. Biol. 47, 2847–2861 (2002).
  Article PubMed CAS Google Scholar
19. Heine, J. J. & Thomas, J. A. Effective X-ray attenuation coefficient measurements from two full field digital mammography systems for data calibration applications. Biomed. Eng. Online 7, 13 (2008) https://doi.org/10.1186/1475-925X-7-13.
20. Fang, Q. et al. Combined optical imaging and mammography of the healthy breast: optical contrast derived from breast structure and compression. IEEE Trans. Med. Imag. 28, 30–42 (2009).
  Article CAS Google Scholar
21. Mallidi, S., Luke, G. P. & Emelianov, S. Photoacoustic imaging in cancer detection, diagnosis, and treatment guidance. Trends Biotechnol. 29, 213–221 (2011).
  Article PubMed PubMed Central CAS Google Scholar
22. Wang, L. V. Multiscale photoacoustic microscopy and computed tomography. Nat. Photon. 3, 503–509 (2009).
  Article ADS CAS Google Scholar
23. Weidner, N., Semple, J. P., Welch, W. R. & Folkman, J. Tumor angiogenesis and metastasis—correlation in invasive breast carcinoma. N. Engl. J. Med. 324, 1–8 (1991).
  Article PubMed CAS Google Scholar
24. Schneider, B. P. & Miller, K. D. Angiogenesis of breast cancer. J. Clin. Oncol. 23, 1782–1790 (2005).
  Article PubMed CAS Google Scholar
25. Reynolds, A. R. et al. Stimulation of tumor growth and angiogenesis by low concetrations of RGD-mimetic integrin inhibitors. Nat. Med. 15, 392–400 (2009).
  Article PubMed CAS Google Scholar
26. Folkman, J. Role of angiogenesis in tumor growth and metastasis. Semin. Oncol. 29, 15–18 (2002).
  Article PubMed CAS Google Scholar
27. Toi, M. et al. Visualization of tumor-related blood vessels in human breast by photoacoustic imaging system with a hemispherical detector array. Sci. Rep. 7, 41970 (2017).
  Article ADS PubMed PubMed Central CAS Google Scholar
28. Kruger, R. A. et al. Dedicated 3D photoacoustic breast imaging. Med. Phys. 40, 113301 (2013).
  Article PubMed PubMed Central Google Scholar
29. Wang, D. et al. Deep tissue photoacoustic computed tomography with fast and compact laser system. Biomed. Opt. Express 8, 112–123 (2017).
  Article PubMed Google Scholar
30. Heijblom, M., Steenbergen, W. & Manohar, S. Clinical photoacoustic breast imaging: the Twente experience. IEEE Pulse 6, 42–46 (2015).
  Article PubMed Google Scholar
31. Heijblom, M. et al. Photoacoustic image patterns of breast carcinoma and comparisons with magnetic resonance imaging and vascular stained histopathlogy. Sci. Rep. 5, 11778 (2015).
  Article ADS PubMed PubMed Central CAS Google Scholar
32. Fakhrejahani, E. et al. Clinical report on the first prototype of a photoacoustic tomography system with dual illumination for breast cancer imaging. PloS ONE 10, e0142287 (2015).
  Article CAS Google Scholar
33. Kitai, T. et al. Photoacoustic mammography: initial clinical results. Breast Cancer 21, 146–153 (2014).
  Article PubMed Google Scholar
34. Ermilov, S. A. et al. Laser optoacoustic imaging system for detection of breast cancer. J. Biomed. Opt. 14, 024007 (2009).
  Article ADS PubMed Google Scholar
35. Ke, H., Erpelding, T. N., Jankovic, L., Liu, C. & Wang, L. V. Performance characterization of an integrated ultrasound, photoacoustic, and thermoacoustic imaging system. J. Biomed. Opt. 17, 056010 (2012).
  Article ADS PubMed PubMed Central CAS Google Scholar
36. Li, X., Heldermon, C. D., Yao, L., Xi, L. & Jiang, H. High resolution functional photoacoustic tomography of breast cancer. Med. Phys. 42, 5321–5328 (2015).
  Article PubMed Google Scholar
37. Cox, B. T., Arridge, S. R. & Beard, P. C. Photoacoustic tomography with a limited-aperture planar sensor and a reverberant cavity. Inverse Probl. 23, S95–S112 (2007).
  Article ADS MathSciNet MATH Google Scholar
38. Huang, B., Xia, J., Maslov, K. & Wang, L. V. Improving limited-view photoacoustic tomography with an acoustic reflector. J. Biomed. Opt. 18, 110505 (2013).
  Article ADS PubMed PubMed Central Google Scholar
39. Li, L. et al. Single-impulse panoramic photoacoustic computed tomography of small animal whole-body dynamics at high spatiotemporal resolution. Nat. BME 1, 0071 (2017).
  CAS Google Scholar
40. Lehman, C. D. & Schnall, M. D. Imaging in breast cancer: magnetic resonance imaing. Breast Cancer Res. 7, 215–219 (2005).
  Article PubMed PubMed Central Google Scholar
41. Smith, A. M., Mancini, M. C. & Nie, S. Bioimaging: second window for in vivo imaging. Nat. Nanotechnol. 4, 710–711 (2009).
  Article ADS PubMed PubMed Central CAS Google Scholar
42. Tsai, P. S. et al. Correlations of neuronal and microvascular densities in murine cortex revealed by direct counting and colocalization of nuclei and vessels. J. Neurosci. 29, 14553–14570 (2009).
  Article PubMed PubMed Central CAS Google Scholar
43. Witt, N. W. et al. A novel measure to characterise optimality of diameter relationships at retinal vascular bifurcations. Artery Res. 4, 75–80 (2010).
  Article PubMed PubMed Central Google Scholar
44. Stanton, A. V. et al. Vascular network changes in the retina with age and hypertension. J. Hypertens. 13, 1724–1728 (1995).
  Article PubMed CAS Google Scholar
45. Witt, N. W. et al. A novel measure to characterise optimality of diameter relationships at retinal vascular bifurcations. Artery Res. 4, 75–80 (2010).
  Article PubMed PubMed Central Google Scholar
46. Davies, P. F. Flow-mediated endothelial mechanotransduction. Physiol. Rev. 75, 519–560 (1995).
  Article PubMed PubMed Central CAS Google Scholar
47. Yoshiya, I., Shimada, Y. & Tanaka, K. Spectrophotometric monitoring of arterial oxygen saturation in the fingertip. Med. Biol. Eng. Comput. 18, 27–32 (1980).
  Article PubMed CAS Google Scholar
48. Xia, J. Calibration-free quantification of absolute oxygen saturation based on the dynamics of photoacoustic signals. Opt. Lett. 38, 2800–2803 (2013).
  Article ADS PubMed PubMed Central CAS Google Scholar
49. Sivaramakrishnan, M. et al. Limitations of quantitative photoacoustic measurements of blood oxygenation in small vessels. Phys. Med. Biol. 52, 1349–1361 (2007).
  Article PubMed Google Scholar
50. Average Breast Size. TheAverageBody.com (2015).
51. Fenner, J. et al. Macroscopic stiffness of breast tumors predicts metastasis. Sci. Rep. 4, 5512 (2014).
  Article PubMed PubMed Central CAS Google Scholar
52. Cancer rate (per 1000 examinations) and Cancer Detection Rate (per 1000 examinations) for 1,838,372 Screening Mammography Examinations from 2004 to 2008 by Age—based on BCSC data through 2009. NCI-funded Breast Cancer Surveillance Consortium (HHSN261201100031C). http://breastscreening.cancer.gov/ (2013).
53. Hendrick, R. E. & Tredennick, T. Benefit to radiation risk of breast-specific gamma imaging compared with mammography in screening asymptomatic women with dense breasts. Radiology 281, 583–588 (2016).
  Article PubMed Google Scholar
54. Jung, H. Assessment of usefulness and risk of mammography screening with exclusive attention to radiation risk. Radiologe 41, 385–395 (2001).
  Article PubMed CAS Google Scholar
55. American National Standards Institute. American National Standard for the Safe Use of Lasers ANSI z136.1–2007 (Laser Institute of America, Orlando, FL, 2007).
56. Caplan, L. R. Carotid-artery disease. N. Engl. J. Med. 315, 886–888 (1986).
  Article PubMed CAS Google Scholar
57. Libby, P., Ridker, P. M. & Maseri, A. Inflammation and atherosclerosis. Circulation 105, 1135–1143 (2002).
  Article PubMed CAS Google Scholar
58. Ouriel, K. Peripheral arterial disease. Lancet 358, 1257–1264 (2005).
  Article Google Scholar
59. Ntziachristos, V., Yodh, A., Schnall, M. D. & Britton, C. MRI-guided diffuse optical spectroscopy of malignant and benign breast lesions. Neoplasia 4, 347–354 (2002).
  Article PubMed PubMed Central Google Scholar
60. Zhu, Q. et al. Benign versus malignant breast masses: optical differentiation with US-guided optical imaging reconstruction. Radiology 237, 57–66 (2005).
  Article PubMed PubMed Central Google Scholar
61. Choe, R. et al. Differentiation of benign and malignant breast tumors by in-vivo three-dimensional parallel-plate diffuse optical tomography. J. Biomed. Opt. 14, 024020 (2009).
  Article ADS PubMed PubMed Central CAS Google Scholar
62. Anastasio, M. A. et al. Half-time image reconstruction in thermoacoustic tomography. IEEE Trans. Med. Imaging 24, 199–210 (2005).
  Article PubMed Google Scholar
63. Yang, X., Li, M. L. & Wang, V. L. Ring-based ultrasonic virtual point detector with applications to photoacoustic tomography. Appl. Phys. Lett. 90, 251103 (2007).
  Article ADS CAS Google Scholar
64. Bender, L. Science Facts: Human Body (Crescent Books, New York, 1992).
65. Lobbes, M. et al. The quality of tumor size assessment by contrast-enhanced spectral mammography and the benefit of additional breast MRI. J. Cancer 6, 144–150 (2015).
  Article PubMed PubMed Central Google Scholar
66. Gruber, I. et al. Measurement of tumor size with mammography, sonography and magnetic resonance imaging as compared to histological tumor size in primary breast cancer. BMC Cancer 13, 328 (2013).
  Article PubMed PubMed Central Google Scholar
67. Manickam, K. et al. Development of a training phantom for compression breast elastography—comparison of various elastography systems and numerical simulations. J. Med. Imaging 2, 047002 (2015).
68. Thirion, J. P. Image matching as a diffusion process: an analogy with Maxwell’s demons. Med. Image Anal. 2, 243–260 (1998).
  Article PubMed CAS Google Scholar
Download references
## Acknowledgements
The authors would like to thank Xi Li and Ruth Holdener for useful discussions and patient recruitment. The authors are grateful to James Linders, Terence T.W. Wong, and Xiaohua Feng for technical support. They also thank Dr. Lily Lai and Professor James Ballard for close reading of the paper. This work was sponsored by the United States National Institutes of Health (NIH) grants DP1 EB016986 (NIH Director’s Pioneer Award), R01 CA186567 (NIH Director’s Transformative Research Award), and R01 EB016963.
## Author information
These authors contributed equally: Li Lin, Peng Hu, Junhui Shi.
Author notes
### Authors and Affiliations
1. Caltech Optical Imaging Laboratory, Andrew and Peggy Cherng Department of Medical Engineering, California Institute of Technology, 1200 East California Boulevard, Pasadena, CA, 91125, USA
  Li Lin, Junhui Shi, Konstantin Maslov, Lei Li & Lihong V. Wang
2. Department of Biomedical Engineering, Washington University in St. Louis, One Brookings Dr., St. Louis, MO, 63130, USA
  Li Lin, Peng Hu & Ruiying Zhang
3. Breast Imaging Section, Washington University School of Medicine in St. Louis, 510 South Kingshighway Blvd, St. Louis, MO, 63108, USA
  Catherine M. Appleton
4. Caltech Optical Imaging Laboratory, Department of Electrical Engineering, California Institute of Technology, 1200 East California Boulevard, Pasadena, CA, 91125, USA
  Lei Li & Lihong V. Wang
### Contributions
L.V.W., K.M., and L.L. conceived and designed the study. L.L., J.S., and K.M. constructed the hardware system. J.S. and L.Li constructed the control program. P.H. developed the software system and the reconstruction algorithm. L.L. and R.Z. performed the experiments. P.H. and L.L. analyzed the data. C.M.A. guided the patient study. L.V.W. supervised the study. All authors wrote the manuscript.
### Corresponding author
Correspondence to Lihong V. Wang.
## Ethics declarations
### Competing interests
L.V.W. has a financial interest in Microphotoacoustics, Inc., CalPACT, LLC, and Union Photoacoustic Technologies, Ltd., which, however, did not support this work. K.M. has a financial interest in Microphotoacoustics, Inc. The remaining authors declare no competing interests.
## Additional information
Publisher's note: Springer Nature remains neutral with regard to jurisdictional claims in published maps and institutional affiliations.
## Electronic supplementary material

### Supplementary Information
### Description of Additional Supplementary Files
### Supplementary Movie 1
### Supplementary Movie 2
### Supplementary Movie 3
### Supplementary Movie 4
## Rights and permissions
Open Access This article is licensed under a Creative Commons Attribution 4.0 International License, which permits use, sharing, adaptation, distribution and reproduction in any medium or format, as long as you give appropriate credit to the original author(s) and the source, provide a link to the Creative Commons license, and indicate if changes were made. The images or other third party material in this article are included in the article’s Creative Commons license, unless indicated otherwise in a credit line to the material. If material is not included in the article’s Creative Commons license and your intended use is not permitted by statutory regulation or exceeds the permitted use, you will need to obtain permission directly from the copyright holder. To view a copy of this license, visit http://creativecommons.org/licenses/by/4.0/.
Reprints and Permissions
# New Method Combines X-rays with Harmless Red Light for Improved Breast Cancer Diagnosis

BY PATRICIA INACIO PHD

IN NEWS.

Click here to subscribe to the Breast Cancer News Newsletter!
A new advance in optical mammography, coupling harmless red or infrared light with X-rays, increases the sensitivity of the approach for breast cancer diagnosis and monitoring by up to 1,000 times, researchers say.
Italian researchers introduced the innovative device at the OSA Biophotonics Congress: Biomedical Optics meeting, held April 3-6 in Hollywood, Florida.
The majority of breast cancer diagnoses are made with X-ray mammography, but this method has limited sensitivity — only 50 to 75 percent of patients with a breast tumor are correctly identified as such — and because it uses ionizing radiation, it’s not without risk. X-ray mammography is also limited by factors such as a patient’s age and body mass index or weight, as well as the use of hormone replacement therapy.
While other imaging techniques, such as ultrasound and magnetic resonance imaging (MRI), are sometimes suggested, none is an effective replacement for X-ray mammography.
Optical mammography is not a new technique. The approach uses red light to measure several parameters, which include blood volume, oxygenation, and the amount of fat, water, and collagen present within the breast. Collagen has been linked to breast cancer onset and progression.
Optical mammography is more appealing than X-rays because it does not use ionizing radiation and does not require strong breast compression. Gentle pressure to the breast tissue is enough, and researchers are now developing 3-D detectors that use no compression at all, instead surrounding the breast tissue with light sources and detectors.
However, a major disadvantage of the technique is its poor spatial resolution, which is necessary to resolve small breast tumors. Researchers hypothesized that combining OM imaging with other imaging methods could help improve this shortcoming.
They developed a new instrument that incorporates red or infrared light with X-rays. The updated instrument includes some changes that eliminate the need for a pre-scan step and has increased sensitivity, making it more robust and cheaper.
Ferocino’s group now plans to assess if optical mammography can be used to monitor breast cancer patients and predict their responses to chemotherapy in clinical studies.
“This technique is able to provide information on the outcome of chemotherapy just weeks after beginning treatment, or possibly even sooner,” co-author of the work Edoardo Ferocino, from the Politecnico di Milano in Italy, said in a press release.
This research is part of a larger European project, called SOLUS, or “Smart Optical and Ultrasound Diagnostics of Breast Cancer,” aimed at combining optical imaging with ultrasound to improve the specificity of breast cancer diagnosis.